## Signals v8
Group | Field Component Vender and Part Number | 2-wire M12 | 4-wire M12 | Field Signal | Location | Version 8 - 2020-06-23 | Units | Eng Min | Eng Max | Raw Min | Raw Max | Wiring Diagram | Mod # | Sig # | Red | Blue | Brown | White | Black | Module | Comments
Relay Outputs |  | 1 |  | 24 VDC | LHS | Pump Enable | Boolean |  |  |  |  | Magnepump v4 | 1 | 0 | [5 empty] | NI 9474
Relay Outputs |  | 2 |  | 24 VDC | LHS | Drive Air Enable | Boolean |  |  |  |  | MAXPRO Filling Cart v4 | 1 | 1 | [5 empty] | NI 9474
Relay Outputs |  | 3 |  | 24 VDC | LHS | Open Vent Valve | Boolean |  |  |  |  | 24 VDC Swagelok Valve v1 | 1 | 2 | [5 empty] | NI 9474 | Convert to single relay?
Relay Outputs |  | 4 |  | 24 VDC | LHS | Close Vent Valve | Boolean |  |  |  |  | 24 VDC Swagelok Valve v1 | 1 | 3 | [5 empty] | NI 9474 | Convert to single relay?
Relay Outputs |  |  |  | 24 VDC |  | Status Alarm | Boolean | [5 empty] | 1 | 4 | [5 empty] | NI 9474
Relay Outputs |  |  |  | 24 VDC |  | Temperature Alert | Boolean | [5 empty] | 1 | 5 | [5 empty] | NI 9474
Relay Outputs |  |  |  | 24 VDC |  | Open Fill Valve | Boolean |  |  |  |  | 24 VDC Swagelok Valve v1 | 1 | 6 | [5 empty] | NI 9474 | Convert to single relay?
Relay Outputs |  |  |  | 24 VDC |  | Close Fill Valve | Boolean |  |  |  |  | 24 VDC Swagelok Valve v1 | 1 | 7 | [5 empty] | NI 9474 | Convert to single relay?
Relay Outputs |  |  |  | 24 VDC |  | Open Pressure Trim Valve | Boolean | [5 empty] | 8 | 0 | [5 empty] | NI 9481 | Convert to single relay?
Relay Outputs |  |  |  | 24 VDC |  | Close Pressure Trim Valve | Boolean | [5 empty] | 8 | 1 | [5 empty] | NI 9481 | Convert to single relay?
Relay Outputs |  |  |  | 24 VDC |  | 1.5 hp Pump Enable | Boolean | [5 empty] | 8 | 2 | [5 empty] | NI 9481
Relay Outputs |  |  |  | 24 VDC | [8 empty] | 8 | 3 | [5 empty] | NI 9481
Cold RTDs | Omega PR-22-3-100-A-1/8-0600-M12 |  |  | 4-wire RTD | RHS | Recup Hot Outlet Temperature [F] | F |  |  |  |  | RTD to NI 9216 | 2 | 0 |  | 2 | 1 | 20 | 21 | NI 9216
Hot RTDs | Conax C16791-1, RTD44W4-INC12-T4(HT)-8.00" |  |  | 4-wire RTD | RHS | Recup Cold Inlet Temperature [F] | F |  |  |  |  | RTD to NI 9216 | 2 | 1 |  | 4 | 3 | 22 | 23 | NI 9216
Hot RTDs | Conax C16791-1, RTD44W4-INC12-T4(HT)-8.00" |  |  | 4-wire RTD | RHS | Heater Outlet Temperature [F] | F |  |  |  |  | RTD to NI 9216 | 2 | 2 |  | 6 | 5 | 24 | 25 | NI 9216
HXer RTDs | Conax C8334-2, RTD43W8-316SS25-T4-8.00" |  |  | 4-wire RTD | RHS | HXer Inlet Temperature [F] | F |  |  |  |  | RTD to NI 9216 | 2 | 3 |  | 8 | 7 | 26 | 27 | NI 9216
HXer RTDs | Conax C8334-2, RTD43W8-316SS25-T4-8.00" |  |  | 4-wire RTD | RHS | HXer Bank 3 West Outlet Temp (B) [F] | F |  |  |  |  | RTD to NI 9216 | 2 | 4 |  | 12 | 11 | 30 | 31 | NI 9216
HXer RTDs | Conax C8334-2, RTD43W8-316SS25-T4-8.00" |  |  | 4-wire RTD | RHS | HXer Bank 3 East Outlet Temp (E) [F] | F |  |  |  |  | RTD to NI 9216 | 2 | 5 |  | 14 | 13 | 32 | 33 | NI 9216
HXer RTDs | Conax C8334-2, RTD43W8-316SS25-T4-8.00" |  |  | 4-wire RTD | RHS | HXer Bank 4 West Outlet Temp (C) [F] | F |  |  |  |  | RTD to NI 9216 | 2 | 6 |  | 16 | 15 | 34 | 35 | NI 9216
HXer RTDs | Conax C8334-2, RTD43W8-316SS25-T4-8.00" |  |  | 4-wire RTD | RHS | HXer Bank 4 East Outlet Temp (F) [F] | F |  |  |  |  | RTD to NI 9216 | 2 | 7 |  | 18 | 17 | 36 | 37 | NI 9216
HXer RTDs | Conax C8334-2, RTD43W8-316SS25-T4-8.00" |  |  | 4-wire RTD | RHS | HXer Bank 2 West Outlet Temp (A) [F] | F |  |  |  |  | RTD to NI 9216 | 3 | 0 |  | 2 | 1 | 20 | 21 | NI 9216 | Change to NI 9216
HXer RTDs | Conax C8334-2, RTD43W8-316SS25-T4-8.00" |  |  | 4-wire RTD | RHS | HXer Bank 2 East Outlet Temp (D) [F] | F |  |  |  |  | RTD to NI 9216 | 3 | 1 |  | 4 | 3 | 22 | 23 | NI 9216 | Change to NI 9216
Cold RTDs | Conax C8334-2, RTD43W8-316SS25-T4-8.00" |  |  | 4-wire RTD | RHS | Cooler Outlet Temperature [F] | F |  |  |  |  | RTD to NI 9219 | 3 | 2 |  | 6 | 5 | 24 | 25 | NI 9216 | Change to NI 9216
Cold RTDs | Omega PR-22-3-100-A-1/8-0600-M12 |  |  | 4-wire RTD | RHS | Cooler Inlet Temperature [F] | F |  |  |  |  | RTD to NI 9219 | 3 | 3 |  | 8 | 7 | 26 | 27 | NI 9216 | Change to NI 9216
Water RTDs | Omega PR-22-3-100-A-1/8-0600-M12 |  |  | 4-wire RTD | RHS | Water Inlet Temperature [F] | F |  |  |  |  | RTD to NI 9219 | 3 | 4 |  | 12 | 11 | 30 | 31 | NI 9216 | Change to NI 9216
Water RTDs | Omega PR-22-3-100-A-1/8-0600-M12 |  |  | 4-wire RTD | RHS | Water Outlet Temperature [F] | F |  |  |  |  | RTD to NI 9219 | 3 | 5 |  | 14 | 13 | 32 | 33 | NI 9216 | Change to NI 9216
Cold RTDs | WIKA |  |  |  | LHS | 1.5 hp Pump Outlet Temperature [F] | F |  |  |  |  | RTD to NI 9219 | 4 | 0 |  | 5 | 3 | 4 | 6 | NI 9219
Hot RTDs | Conax C17049-1, RTD44W4-INC25-T4(HT)-8.00" |  | 14 | Tri-Loop F #2 | RHS | HXer Outlet Temperature [F] | F |  |  |  |  | RTD to NI 9219 | 4 | 1 |  | 5 | 3 | 4 | 6 | NI 9219
[6 empty] | 1.5 hp Pump dP [psid] | psid | 0 | 27 | 0.004 | 0.02 | Kobold Pressure v3 | 4 | 2 |  |  | Vsup | 3 |  | NI 9219
[6 empty] | 1.5 hp Pump Frequency [Hz] | Hz | 0 | 76.5 | 0.004 | 0.02 | 1.5 hp Magnepump v1 | 4 | 3 |  |  | 3 | 5 |  | NI 9219
mA Outputs |  |  | 1 | 4-20 mA | LHS | Pump Frequency Cmd [Hz] | Hz | 0 | 76.5 | 0.004 | 0.02 | Magnepump v4 | 5 | 0 |  | AO0 |  |  | COM0 | NI 9266
mA Outputs |  |  | 2 | 4-20 mA | LHS | Heater Power Cmd [%] | % | 0 | 100 | 0.004 | 0.02 | KAPL Heater Controller v2 | 5 | 1 |  | AO1 |  |  | COM1 | NI 9266
mA Outputs |  |  | 3 | 4-20 mA | LHS | Recirc Valve % Closed Cmd | % | 0 | 100 | 0.004 | 0.01488 | TRIAC Valve v3 | 5 | 2 |  | AO2 |  |  | COM2 | NI 9266
mA Outputs |  |  | 4 | 4-20 mA | LHS | Flow Valve % Closed Cmd | % | 0 | 100 | 0.004 | 0.02 | TRIAC Valve v3 | 5 | 3 |  | AO3 |  |  | COM3 | NI 9266
mA Outputs |  |  | 5 | 4-20 mA | LHS | Recup Bypass Valve % Closed Cmd | % | 0 | 100 | 0.004 | 0.02 | TRIAC Valve v3 | 5 | 4 |  | AO4 |  |  | COM4 | NI 9266
mA Outputs |  |  | 6 | 4-20 mA | LHS | Drive Air Pressure Cmd | psig | 0 | 141.2708 | 0.004 | 0.02 | MAXPRO Filling Cart v4 | 5 | 5 |  | COM5 | Vsup | AO5 |  | NI 9266
mA Outputs |  |  | 7 | 4-20 mA | LHS | Mixing Valve % Closed Cmd | % | 0 | 100 | 0.004 | 0.02 |  | 5 | 6 |  | AO6 |  |  | COM6 | NI 9266
mA Outputs |  |  | 8 | 4-20 mA | LHS | 1.5 hp Pump Frequency Cmd [Hz] | Hz | 0 | 76.5 | 0.004 | 0.02 | 1.5 hp Magnepump v1 | 5 | 7 |  | AO7 |  |  | COM | NI 9266
Water Flow |  |  | 17 | 4-20 mA | LHS | Water Flow Rate [gpm] | gpm | 0 | 32 | 0.004 | 0.02 | Water Flow Meter v1 | 6 | 0 |  |  |  | 20 | 1 | NI 9208
Pressures | Rosemount |  | 10 | Tri-Loop B BURST | LHS | Cooler CO2 Inlet Pressure [psig] | psig | 0 | 3626 | 0.004 | 0.02 | Rosemount v4 | 6 | 1 | 21 |  |  |  | 2 | NI 9208
Pressure Drops | Rosemount |  | 10 | Tri-Loop B #1 | LHS | Cooler CO2 dP [psid] | psid | 0 | 36.1273 | 0.004 | 0.02 | Rosemount v4 | 6 | 2 | 22 |  |  |  | 3 | NI 9208
Cold RTDs | Omega PR-22-3-100-A-1/8-0600-M12 |  | 10 | Tri-Loop B #2 | LHS | Pump Outlet Temperature [F] | F | -328 | 1562 | 0.004 | 0.02 | Rosemount v4 | 6 | 3 | 23 |  |  |  | 4 | NI 9208
Feedback Signals |  |  | 1 | 4-20 mA | LHS | Pump Frequency [Hz] | Hz | 0 | 76.5 | 0.004 | 0.02 | Magnepump v4 | 6 | 4 |  |  | 5 | 29 |  | NI 9208
Feedback Signals |  |  | 2 | 4-20 mA | LHS | Heater Power [%] | % | 0 | 100 | 0.004 | 0.02 | KAPL Heater Controller v2 | 6 | 5 |  |  | 6 | 29 |  | NI 9208
Pressures | Rosemount 3051SMV3P14A4R3A11A1JL2Q4GMA1039 |  | 9 | Tri-Loop A BURST | LHS | Pump Outlet Pressure [psig] | psig | 0 | 3626 | 0.004 | 0.02 | Rosemount v4 | 6 | 6 | 26 |  |  |  | 7 | NI 9208
Pressure Drops | Rosemount 3051SMV3P14A4R3A11A1JL2Q4GMA1039 |  | 9 | Tri-Loop A #1 | LHS | Pump dP [psid] | psid | 0 | 150 | 0.004 | 0.02 | Rosemount v4 | 6 | 7 | 27 |  |  |  | 8 | NI 9208
Cold RTDs | Conax C8334-2, RTD43W8-316SS25-T4-8.00" |  | 9 | Tri-Loop A #2 | LHS | Pump Inlet Temperature [F] | F | -328 | 1562 | 0.004 | 0.02 | Rosemount v4 | 6 | 8 | 30 |  |  |  | 11 | NI 9208
Feedback Signals |  |  | 3 | 4-20 mA | LHS | Recirc Valve % Closed | % | 0 | 100 | 0.004 | 0.02 | TRIAC Valve v3 | 6 | 9 |  |  | 12 | 9 |  | NI 9208
Feedback Signals |  |  | 4 | 4-20 mA | LHS | Flow Valve % Closed | % | 0 | 100 | 0.004 | 0.02 | TRIAC Valve v3 | 6 | 10 |  |  | 13 | 9 |  | NI 9208
Feedback Signals |  |  | 5 | 4-20 mA | LHS | Recup Bypass Valve % Closed | % | 0 | 100 | 0.004 | 0.02 | TRIAC Valve v3 | 6 | 11 |  |  | 14 | 9 |  | NI 9208
Flows | Emerson CMFS100P999N2BAEKZX_22757 |  | 16 | 4-20 mA | LHS | Pump Flow Rate [kg¦s] | kg/s | 0 | 6.5277 | 0.004 | 0.02 | Coriolis Flowmeter v3 | 6 | 12 |  | 15 |  |  | 28 | NI 9208
Densities | Emerson CMFS100P999N2BAEKZX_22757 |  | 16 | 4-20 mA | LHS | Pump Outlet Density [kg¦m3] | kg/m^3 | 0 | 5000 | 0.004 | 0.02 | Coriolis Flowmeter v3 | 6 | 13 |  |  | 16 | 28 |  | NI 9208
Water Pressures |  |  | 18 | 4-20 mA | LHS | Water Inlet Pressure [psig] | psig | 0 | 140 | 0.004 | 0.02 | Water Pressure Transducer v1 | 6 | 14 |  | 10 | 36 | 17 |  | NI 9208
Water Pressures |  |  | 19 | 4-20 mA | LHS | Water Outlet Pressure [psig] | psig | 0 | 140 | 0.004 | 0.02 | Water Pressure Transducer v1 | 6 | 15 |  | 10 | 37 | 18 |  | NI 9208
Pressure Drops | Rosemount 3051SMV3P14A4R3A11A1JL2Q4GMA1039 |  | 11 | Tri-Loop C #1 | LHS | Recup Cold dP [psid] | psid | 0 | 150 | 0.004 | 0.02 | Rosemount v4 | 7 | 0 | 20 |  |  |  | 1 | NI 9208
Cold RTDs | Omega PR-22-3-100-A-1/8-0600-M12 |  | 11 | Tri-Loop C #2 | LHS | Recup Cold Outlet Temperature [F] | F | -328 | 1562 | 0.004 | 0.02 | Rosemount v4 | 7 | 1 | 21 |  |  |  | 2 | NI 9208
Pressures | Rosemount 3051SMV3P14A4R3A11A1JL2Q4GMA1039 |  | 11 | Tri-Loop C BURST | LHS | Recup Cold Inlet Pressure [psig] | psig | 0 | 3626 | 0.004 | 0.02 | Rosemount v4 | 7 | 2 | 22 |  |  |  | 3 | NI 9208
Pressure Drops | Rosemount 3051SMV3P14A4R3A11A1JL2Q4GMA1039 |  | 12 | Tri-Loop D #1 | RHS | Recup Hot dP [psid] | psid | 0 | 150 | 0.004 | 0.02 | Rosemount v4 | 7 | 3 | 23 |  |  |  | 4 | NI 9208
Hot RTDs | Conax C16791-1, RTD44W4-INC12-T4(HT)-8.00" |  | 12 | Tri-Loop D #2 | RHS | Recup Hot Inlet Temperature [F] | F | -328 | 1562 | 0.004 | 0.02 | Rosemount v4 | 7 | 4 | 24 |  |  |  | 5 | NI 9208
Pressures | Rosemount 3051SMV3P14A4R3A11A1JL2Q4GMA1039 |  | 12 | Tri-Loop D BURST | RHS | Recup Hot Inlet Pressure [psig] | psig | 0 | 3626 | 0.004 | 0.02 | Rosemount v4 | 7 | 5 | 25 |  |  |  | 6 | NI 9208
Flows | Emerson CMFS100P999N2BAEKZX_22757 |  | 15 | 4-20 mA | LHS | Recup Cold Inlet Flow Rate [kg¦s] | kg/s | 0 | 6.5277 | 0.004 | 0.02 | Coriolis Flowmeter v3 | 7 | 6 |  | 7 |  |  | 9 | NI 9208
Densities | Emerson CMFS100P999N2BAEKZX_22757 |  | 15 | 4-20 mA | LHS | Recup Cold Inlet Density [kg¦m^3] | kg/m^3 | 0 | 5000 | 0.004 | 0.02 | Coriolis Flowmeter v3 | 7 | 7 |  |  | 8 | 9 |  | NI 9208
Pressure Drops | Rosemount 3051SMV3P14A4R3A11A1JL2Q4GMA1039 |  | 14 | Tri-Loop F #1 | RHS | HXer dP [psid] | psid | 0 | 150 | 0.004 | 0.02 | Rosemount v4 | 7 | 8 | 30 |  |  |  | 11 | NI 9208
Pressures | WIKA |  | 22 | 4-20 mA | LHS | Regulated Pressure [psig] | psig | 0 | 3000 | 0.004 | 0.02 | WIKA Pressure v1 | 7 | 9 | 31 |  |  |  | 12 | NI 9208
Pressures | Rosemount 3051SMV3P14A4R3A11A1JL2Q4GMA1039 |  | 14 | Tri-Loop F BURST | RHS | HXer Inlet Pressure [psig] | psi? | 0 | 3626 | 0.004 | 0.02 | Rosemount v4 | 7 | 10 | 32 |  |  |  | 13 | NI 9208
Pressures | WIKA |  | 21 | 4-20 mA | LHS | Supply Pressure [psig] | psig | 0 | 3000 | 0.004 | 0.02 | WIKA Pressure v1 | 7 | 11 | 33 |  |  |  | 14 | NI 9208
Pressure Drops | Rosemount |  | 13 | Tri-Loop E #1 | RHS | Heater dP [psid] | psid | 0 | 36.1273 | 0.004 | 0.02 | Rosemount v4 | 7 | 12 | 34 |  |  |  | 15 | NI 9208
HXer RTDs | Conax C16791-1, RTD44W4-INC12-T4(HT)-8.00" |  | 13 | Tri-Loop E #2 | RHS | Heater Inlet Temperature [F] | F | -328 | 1562 | 0.004 | 0.02 | Rosemount v4 | 7 | 13 | 35 |  |  |  | 16 | NI 9208
Pressures | Rosemount |  | 13 | Tri-Loop E BURST | RHS | Heater Inlet Pressure [psig] | psig | 0 | 3626 | 0.004 | 0.02 | Rosemount v4 | 7 | 14 | 36 |  |  |  | 17 | NI 9208
Feedback Signals |  |  | 7 | 4-20 mA | LHS | Mixing Valve % Closed | % | 0 | 100 | 0.004 | 0.02 |  | 7 | 15 |  |  | 18 | 9, 10, 28, or 29 |  | NI 9208
|  |  |  | K-type TC | Bottom | Vent Temperature [F] | F | [5 empty] | 8 | [6 empty] | NI 9213
|  |  |  | K-type TC | Bottom | Supply Temperature [F] | F | [5 empty] | 8 | [6 empty] | NI 9213
|  |  |  | K-type TC | Bottom | Regulated Temperature [F] | F | [5 empty] | 8 | [6 empty] | NI 9213
|  |  |  | K-type TC | Bottom | Compressor Temperature [F] | F | [5 empty] | 8 | [6 empty] | NI 9213
|  |  |  | K-type TC | Bottom | KAPL Heater Lower RHS [F] | F | [5 empty] | 8 | [6 empty] | NI 9213
|  |  |  | K-type TC | Bottom | KAPL Heater Lower LHS [F] | F | [5 empty] | 8 | [6 empty] | NI 9213
|  |  |  | K-type TC | Bottom | KAPL Heater Upper RHS [F] | F | [5 empty] | 8 | [6 empty] | NI 9213
|  |  |  | K-type TC | Bottom | KAPL Heater Upper LHS [F] | F | [5 empty] | 8 | [6 empty] | NI 9213
|  |  |  | K-type TC | Bottom | Solex HXer? | F | [5 empty] | 8 | [6 empty] | NI 9213
|  |  |  | K-type TC | Bottom | Solex HXer? | F | [5 empty] | 8 | [6 empty] | NI 9213
|  |  |  | K-type TC | Bottom | Solex HXer? | F | [5 empty] | 8 | [6 empty] | NI 9213
|  |  |  | K-type TC | Bottom | Solex HXer? | F | [5 empty] | 8 | [6 empty] | NI 9213
|  |  |  | K-type TC | Bottom | Solex HXer? | F | [5 empty] | 8 | [6 empty] | NI 9213
|  |  |  | K-type TC | Bottom | Solex HXer? | F | [5 empty] | 8 | [6 empty] | NI 9213
|  |  |  | K-type TC | Bottom | Solex HXer? | F | [5 empty] | 8 | [6 empty] | NI 9213
|  |  |  | K-type TC | Bottom | Solex HXer? | F | [5 empty] | 8 | [6 empty] | NI 9213

## Signals v7
Group | Field Component Vender and Part Number | 2-wire M12 | 4-wire M12 | Field Signal | Location | Version 5 - 2019-05-22 | Units | Eng Min | Eng Max | Raw Min | Raw Max | Wiring Diagram | Mod # | Sig # | Red | Blue | Brown | White | Black | Module | Comments
Relay Outputs |  | 1 |  | 24 VDC | LHS | Pump Enable | Boolean |  |  |  |  | Magnepump v4 | 1 | 0 | [5 empty] | NI 9474
Relay Outputs |  | 2 |  | 24 VDC | LHS | Drive Air Enable | Boolean |  |  |  |  | MAXPRO Filling Cart v4 | 1 | 1 | [5 empty] | NI 9474
Relay Outputs |  | 3 |  | 24 VDC | LHS | Open Vent Valve | Boolean |  |  |  |  | 24 VDC Swagelok Valve v1 | 1 | 2 | [5 empty] | NI 9474 | Convert to single relay?
Relay Outputs |  | 4 |  | 24 VDC | LHS | Close Vent Valve | Boolean |  |  |  |  | 24 VDC Swagelok Valve v1 | 1 | 3 | [5 empty] | NI 9474 | Convert to single relay?
Relay Outputs |  |  |  | 24 VDC |  | Status Alarm | Boolean | [5 empty] | 1 | 4 | [5 empty] | NI 9474
Relay Outputs |  |  |  | 24 VDC |  | Temperature Alert | Boolean | [5 empty] | 1 | 5 | [5 empty] | NI 9474
Relay Outputs |  |  |  | 24 VDC |  | Open Fill Valve | Boolean |  |  |  |  | 24 VDC Swagelok Valve v1 | 1 | 6 | [5 empty] | NI 9474 | Convert to single relay?
Relay Outputs |  |  |  | 24 VDC |  | Close Fill Valve | Boolean |  |  |  |  | 24 VDC Swagelok Valve v1 | 1 | 7 | [5 empty] | NI 9474 | Convert to single relay?
Relay Outputs |  |  |  | 24 VDC |  | Open Pressure Trim Valve | Boolean | [5 empty] | 8 | 0 | [5 empty] | NI 9481 | Convert to single relay?
Relay Outputs |  |  |  | 24 VDC |  | Close Pressure Trim Valve | Boolean | [5 empty] | 8 | 1 | [5 empty] | NI 9481 | Convert to single relay?
Relay Outputs |  |  |  | 24 VDC |  | 1.5 hp Pump Enable | Boolean | [5 empty] | 8 | 2 | [5 empty] | NI 9481
Relay Outputs |  |  |  | 24 VDC | [8 empty] | 8 | 3 | [5 empty] | NI 9481
Cold RTDs | Omega PR-22-3-100-A-1/8-0600-M12 |  |  | 4-wire RTD | RHS | Recup Hot Outlet Temperature [F] | F |  |  |  |  | RTD to NI 9216 | 2 | 0 |  | 2 | 1 | 20 | 21 | NI 9216
Hot RTDs | Conax C16791-1, RTD44W4-INC12-T4(HT)-8.00" |  |  | 4-wire RTD | RHS | Recup Cold Inlet Temperature [F] | F |  |  |  |  | RTD to NI 9216 | 2 | 1 |  | 4 | 3 | 22 | 23 | NI 9216
Hot RTDs | Conax C16791-1, RTD44W4-INC12-T4(HT)-8.00" |  |  | 4-wire RTD | RHS | Heater Outlet Temperature [F] | F |  |  |  |  | RTD to NI 9216 | 2 | 2 |  | 6 | 5 | 24 | 25 | NI 9216
HXer RTDs | Conax C8334-2, RTD43W8-316SS25-T4-8.00" |  |  | 4-wire RTD | RHS | HXer Inlet Temperature [F] | F |  |  |  |  | RTD to NI 9216 | 2 | 3 |  | 8 | 7 | 26 | 27 | NI 9216
HXer RTDs | Conax C8334-2, RTD43W8-316SS25-T4-8.00" |  |  | 4-wire RTD | RHS | HXer Bank 3 West Outlet Temp (B) [F] | F |  |  |  |  | RTD to NI 9216 | 2 | 4 |  | 12 | 11 | 30 | 31 | NI 9216
HXer RTDs | Conax C8334-2, RTD43W8-316SS25-T4-8.00" |  |  | 4-wire RTD | RHS | HXer Bank 3 East Outlet Temp (E) [F] | F |  |  |  |  | RTD to NI 9216 | 2 | 5 |  | 14 | 13 | 32 | 33 | NI 9216
HXer RTDs | Conax C8334-2, RTD43W8-316SS25-T4-8.00" |  |  | 4-wire RTD | RHS | HXer Bank 4 West Outlet Temp (C) [F] | F |  |  |  |  | RTD to NI 9216 | 2 | 6 |  | 16 | 15 | 34 | 35 | NI 9216
HXer RTDs | Conax C8334-2, RTD43W8-316SS25-T4-8.00" |  |  | 4-wire RTD | RHS | HXer Bank 4 East Outlet Temp (F) [F] | F |  |  |  |  | RTD to NI 9216 | 2 | 7 |  | 18 | 17 | 36 | 37 | NI 9216
HXer RTDs | Conax C8334-2, RTD43W8-316SS25-T4-8.00" |  |  | 4-wire RTD | RHS | HXer Bank 2 West Outlet Temp (A) [F] | F |  |  |  |  | RTD to NI 9216 | [7 empty] | NI 9216
HXer RTDs | Conax C8334-2, RTD43W8-316SS25-T4-8.00" |  |  | 4-wire RTD | RHS | HXer Bank 2 East Outlet Temp (D) [F] | F |  |  |  |  | RTD to NI 9216 | [7 empty] | NI 9216
| Conax C8334-2, RTD43W8-316SS25-T4-8.00" |  |  | 4-wire RTD | RHS |  | F |  |  |  |  | RTD to NI 9216 | [7 empty] | NI 9216
|  |  |  | 4-wire RTD | RHS |  | F |  |  |  |  | RTD to NI 9216 | [7 empty] | NI 9216
Cold RTDs | Conax C8334-2, RTD43W8-316SS25-T4-8.00" |  |  | 4-wire RTD | RHS | Cooler Outlet Temperature [F] | F |  |  |  |  | RTD to NI 9219 | 3 | 0 |  | 5 | 3 | 4 | 6 | NI 9219 | Change to NI 9216?
Cold RTDs | Omega PR-22-3-100-A-1/8-0600-M12 |  |  | 4-wire RTD | RHS | Cooler Inlet Temperature [F] | F |  |  |  |  | RTD to NI 9219 | 3 | 1 |  | 5 | 3 | 4 | 6 | NI 9219 | Change to NI 9216?
Water RTDs | Omega PR-22-3-100-A-1/8-0600-M12 |  |  | 4-wire RTD | RHS | Water Inlet Temperature [F] | F |  |  |  |  | RTD to NI 9219 | 3 | 2 |  | 5 | 3 | 4 | 6 | NI 9219 | Change to NI 9216?
Water RTDs | Omega PR-22-3-100-A-1/8-0600-M12 |  |  | 4-wire RTD | RHS | Water Outlet Temperature [F] | F |  |  |  |  | RTD to NI 9219 | 3 | 3 |  | 5 | 3 | 4 | 6 | NI 9219 | Change to NI 9216?
Cold RTDs | WIKA |  |  |  | LHS | 1.5 hp Pump Outlet Temperature [F] | F |  |  |  |  | RTD to NI 9219 | 4 | 0 |  | 5 | 3 | 4 | 6 | NI 9219
Hot RTDs | Conax C17049-1, RTD44W4-INC25-T4(HT)-8.00" |  | 14 | Tri-Loop F #2 | RHS | HXer Outlet Temperature [F] | F |  |  |  |  | RTD to NI 9219 | 4 | 1 |  | 5 | 3 | 4 | 6 | NI 9219
[6 empty] | 1.5 hp Pump dP [psid] | psid | 0 | 27 | 0.004 | 0.02 | Kobold Pressure v3 | 4 | 2 |  |  | Vsup | 3 |  | NI 9219
[6 empty] | 1.5 hp Pump Frequency [Hz] | Hz | 0 | 76.5 | 0.004 | 0.02 | 1.5 hp Magnepump v1 | 4 | 3 |  |  | 3 | 5 |  | NI 9219
mA Outputs |  |  | 1 | 4-20 mA | LHS | Pump Frequency Cmd [Hz] | Hz | 0 | 76.5 | 0.004 | 0.02 | Magnepump v4 | 5 | 0 |  | AO0 |  |  | COM0 | NI 9266
mA Outputs |  |  | 2 | 4-20 mA | LHS | Heater Power Cmd [%] | % | 0 | 100 | 0.004 | 0.02 | KAPL Heater Controller v2 | 5 | 1 |  | AO1 |  |  | COM1 | NI 9266
mA Outputs |  |  | 3 | 4-20 mA | LHS | Recirc Valve % Closed Cmd | % | 0 | 100 | 0.004 | 0.01488 | TRIAC Valve v3 | 5 | 2 |  | AO2 |  |  | COM2 | NI 9266
mA Outputs |  |  | 4 | 4-20 mA | LHS | Flow Valve % Closed Cmd | % | 0 | 100 | 0.004 | 0.02 | TRIAC Valve v3 | 5 | 3 |  | AO3 |  |  | COM3 | NI 9266
mA Outputs |  |  | 5 | 4-20 mA | LHS | Recup Bypass Valve % Closed Cmd | % | 0 | 100 | 0.004 | 0.02 | TRIAC Valve v3 | 5 | 4 |  | AO4 |  |  | COM4 | NI 9266
mA Outputs |  |  | 6 | 4-20 mA | LHS | Drive Air Pressure Cmd | psig | 0 | 141.2708 | 0.004 | 0.02 | MAXPRO Filling Cart v4 | 5 | 5 |  | COM5 | Vsup | AO5 |  | NI 9266
mA Outputs |  |  | 7 | 4-20 mA | LHS | Water Valve % Closed Cmd | % | 0 | 100 | 0.004 | 0.02 |  | 5 | 6 | [5 empty] | NI 9266
mA Outputs |  |  | 8 | 4-20 mA | LHS | 1.5 hp Pump Frequency Cmd [Hz] | Hz | 0 | 76.5 | 0.004 | 0.02 | 1.5 hp Magnepump v1 | 5 | 7 |  | DO7 |  |  | COM | NI 9266
Water Flow |  |  | 17 | 4-20 mA | LHS | Water Flow Rate [gpm] | gpm | 0 | 32 | 0.004 | 0.02 | Water Flow Meter v1 | 6 | 0 |  |  |  | 20 | 1 | NI 9208
Pressures | Rosemount |  | 10 | Tri-Loop B BURST | LHS | Cooler CO2 Inlet Pressure [psig] | psig | 0 | 3626 | 0.004 | 0.02 | Rosemount v4 | 6 | 1 | 21 |  |  |  | 2 | NI 9208
Pressure Drops | Rosemount |  | 10 | Tri-Loop B #1 | LHS | Cooler CO2 dP [psid] | psid | 0 | 36.1273 | 0.004 | 0.02 | Rosemount v4 | 6 | 2 | 22 |  |  |  | 3 | NI 9208
Cold RTDs | Omega PR-22-3-100-A-1/8-0600-M12 |  | 10 | Tri-Loop B #2 | LHS | Pump Outlet Temperature [F] | F | -328 | 1562 | 0.004 | 0.02 | Rosemount v4 | 6 | 3 | 23 |  |  |  | 4 | NI 9208
Feedback Signals |  |  | 1 | 4-20 mA | LHS | Pump Frequency [Hz] | Hz | 0 | 76.5 | 0.004 | 0.02 | Magnepump v4 | 6 | 4 |  |  | 5 | 29 |  | NI 9208
Feedback Signals |  |  | 2 | 4-20 mA | LHS | Heater Power [%] | % | 0 | 100 | 0.004 | 0.02 | KAPL Heater Controller v2 | 6 | 5 |  |  | 6 | 29 |  | NI 9208
Pressures | Rosemount 3051SMV3P14A4R3A11A1JL2Q4GMA1039 |  | 9 | Tri-Loop A BURST | LHS | Pump Outlet Pressure [psig] | psig | 0 | 3626 | 0.004 | 0.02 | Rosemount v4 | 6 | 6 | 26 |  |  |  | 7 | NI 9208
Pressure Drops | Rosemount 3051SMV3P14A4R3A11A1JL2Q4GMA1039 |  | 9 | Tri-Loop A #1 | LHS | Pump dP [psid] | psid | 0 | 150 | 0.004 | 0.02 | Rosemount v4 | 6 | 7 | 27 |  |  |  | 8 | NI 9208
Cold RTDs | Conax C8334-2, RTD43W8-316SS25-T4-8.00" |  | 9 | Tri-Loop A #2 | LHS | Pump Inlet Temperature [F] | F | -328 | 1562 | 0.004 | 0.02 | Rosemount v4 | 6 | 8 | 30 |  |  |  | 11 | NI 9208
Feedback Signals |  |  | 3 | 4-20 mA | LHS | Recirc Valve % Closed | % | 0 | 100 | 0.004 | 0.02 | TRIAC Valve v3 | 6 | 9 |  |  | 12 | 9 |  | NI 9208
Feedback Signals |  |  | 4 | 4-20 mA | LHS | Flow Valve % Closed | % | 0 | 100 | 0.004 | 0.02 | TRIAC Valve v3 | 6 | 10 |  |  | 13 | 9 |  | NI 9208
Feedback Signals |  |  | 5 | 4-20 mA | LHS | Recup Bypass Valve % Closed | % | 0 | 100 | 0.004 | 0.02 | TRIAC Valve v3 | 6 | 11 |  |  | 14 | 9 |  | NI 9208
Flows | Emerson CMFS100P999N2BAEKZX_22757 |  | 16 | 4-20 mA | LHS | Pump Flow Rate [kg¦s] | kg/s | 0 | 6.5277 | 0.004 | 0.02 | Coriolis Flowmeter v3 | 6 | 12 |  | 15 |  |  | 28 | NI 9208
Densities | Emerson CMFS100P999N2BAEKZX_22757 |  | 16 | 4-20 mA | LHS | Pump Outlet Density [kg¦m3] | kg/m^3 | 0 | 5000 | 0.004 | 0.02 | Coriolis Flowmeter v3 | 6 | 13 |  |  | 16 | 28 |  | NI 9208
Water Pressures |  |  | 18 | 4-20 mA | LHS | Water Inlet Pressure [psig] | psig | 0 | 140 | 0.004 | 0.02 | Water Pressure Transducer v1 | 6 | 14 |  | 10 | 36 | 17 |  | NI 9208
Water Pressures |  |  | 19 | 4-20 mA | LHS | Water Outlet Pressure [psig] | psig | 0 | 140 | 0.004 | 0.02 | Water Pressure Transducer v1 | 6 | 15 |  | 10 | 37 | 18 |  | NI 9208
Pressure Drops | Rosemount 3051SMV3P14A4R3A11A1JL2Q4GMA1039 |  | 11 | Tri-Loop C #1 | LHS | Recup Cold dP [psid] | psid | 0 | 150 | 0.004 | 0.02 | Rosemount v4 | 7 | 0 | 20 |  |  |  | 1 | NI 9208
Cold RTDs | Omega PR-22-3-100-A-1/8-0600-M12 |  | 11 | Tri-Loop C #2 | LHS | Recup Cold Outlet Temperature [F] | F | -328 | 1562 | 0.004 | 0.02 | Rosemount v4 | 7 | 1 | 21 |  |  |  | 2 | NI 9208
Pressures | Rosemount 3051SMV3P14A4R3A11A1JL2Q4GMA1039 |  | 11 | Tri-Loop C BURST | LHS | Recup Cold Inlet Pressure [psig] | psig | 0 | 3626 | 0.004 | 0.02 | Rosemount v4 | 7 | 2 | 22 |  |  |  | 3 | NI 9208
Pressure Drops | Rosemount 3051SMV3P14A4R3A11A1JL2Q4GMA1039 |  | 12 | Tri-Loop D #1 | RHS | Recup Hot dP [psid] | psid | 0 | 150 | 0.004 | 0.02 | Rosemount v4 | 7 | 3 | 23 |  |  |  | 4 | NI 9208
Hot RTDs | Conax C16791-1, RTD44W4-INC12-T4(HT)-8.00" |  | 12 | Tri-Loop D #2 | RHS | Recup Hot Inlet Temperature [F] | F | -328 | 1562 | 0.004 | 0.02 | Rosemount v4 | 7 | 4 | 24 |  |  |  | 5 | NI 9208
Pressures | Rosemount 3051SMV3P14A4R3A11A1JL2Q4GMA1039 |  | 12 | Tri-Loop D BURST | RHS | Recup Hot Inlet Pressure [psig] | psig | 0 | 3626 | 0.004 | 0.02 | Rosemount v4 | 7 | 5 | 25 |  |  |  | 6 | NI 9208
Flows | Emerson CMFS100P999N2BAEKZX_22757 |  | 15 | 4-20 mA | LHS | Recup Cold Inlet Flow Rate [kg¦s] | kg/s | 0 | 6.5277 | 0.004 | 0.02 | Coriolis Flowmeter v3 | 7 | 6 |  | 7 |  |  | 9 | NI 9208
Densities | Emerson CMFS100P999N2BAEKZX_22757 |  | 15 | 4-20 mA | LHS | Recup Cold Inlet Density [kg¦m^3] | kg/m^3 | 0 | 5000 | 0.004 | 0.02 | Coriolis Flowmeter v3 | 7 | 7 |  |  | 8 | 9 |  | NI 9208
Pressure Drops | Rosemount 3051SMV3P14A4R3A11A1JL2Q4GMA1039 |  | 14 | Tri-Loop F #1 | RHS | HXer dP [psid] | psid | 0 | 150 | 0.004 | 0.02 | Rosemount v4 | 7 | 8 | 30 |  |  |  | 11 | NI 9208
Pressures | WIKA |  | 22 | 4-20 mA | LHS | Regulated Pressure [psig] | psig | 0 | 3000 | 0.004 | 0.02 | WIKA Pressure v1 | 7 | 9 | 31 |  |  |  | 12 | NI 9208
Pressures | Rosemount 3051SMV3P14A4R3A11A1JL2Q4GMA1039 |  | 14 | Tri-Loop F BURST | RHS | HXer Inlet Pressure [psig] | psi? | 0 | 3626 | 0.004 | 0.02 | Rosemount v4 | 7 | 10 | 32 |  |  |  | 13 | NI 9208
Pressures | WIKA |  | 21 | 4-20 mA | LHS | Supply Pressure [psig] | psig | 0 | 3000 | 0.004 | 0.02 | WIKA Pressure v1 | 7 | 11 | 33 |  |  |  | 14 | NI 9208
Pressure Drops | Rosemount |  | 13 | Tri-Loop E #1 | RHS | Heater dP [psid] | psid | 0 | 36.1273 | 0.004 | 0.02 | Rosemount v4 | 7 | 12 | 34 |  |  |  | 15 | NI 9208
HXer RTDs | Conax C16791-1, RTD44W4-INC12-T4(HT)-8.00" |  | 13 | Tri-Loop E #2 | RHS | Heater Inlet Temperature [F] | F | -328 | 1562 | 0.004 | 0.02 | Rosemount v4 | 7 | 13 | 35 |  |  |  | 16 | NI 9208
Pressures | Rosemount |  | 13 | Tri-Loop E BURST | RHS | Heater Inlet Pressure [psig] | psig | 0 | 3626 | 0.004 | 0.02 | Rosemount v4 | 7 | 14 | 36 |  |  |  | 17 | NI 9208
Feedback Signals |  |  | 7 | 4-20 mA | LHS | Water Valve % Closed | % | 0 | 100 | 0.004 | 0.02 |  | 7 | 15 | [5 empty] | NI 9208
|  |  |  | K-type TC | Bottom | Vent Temperature [F] | F | [5 empty] | 8 | [6 empty] | NI 9213
|  |  |  | K-type TC | Bottom | Supply Temperature [F] | F | [5 empty] | 8 | [6 empty] | NI 9213
|  |  |  | K-type TC | Bottom | Regulated Temperature [F] | F | [5 empty] | 8 | [6 empty] | NI 9213
|  |  |  | K-type TC | Bottom | Compressor Temperature [F] | F | [5 empty] | 8 | [6 empty] | NI 9213
|  |  |  | K-type TC | Bottom | KAPL Heater Lower RHS [F] | F | [5 empty] | 8 | [6 empty] | NI 9213
|  |  |  | K-type TC | Bottom | KAPL Heater Lower LHS [F] | F | [5 empty] | 8 | [6 empty] | NI 9213
|  |  |  | K-type TC | Bottom | KAPL Heater Upper RHS [F] | F | [5 empty] | 8 | [6 empty] | NI 9213
|  |  |  | K-type TC | Bottom | KAPL Heater Upper LHS [F] | F | [5 empty] | 8 | [6 empty] | NI 9213
|  |  |  | K-type TC | Bottom | Solex HXer? | F | [5 empty] | 8 | [6 empty] | NI 9213
|  |  |  | K-type TC | Bottom | Solex HXer? | F | [5 empty] | 8 | [6 empty] | NI 9213
|  |  |  | K-type TC | Bottom | Solex HXer? | F | [5 empty] | 8 | [6 empty] | NI 9213
|  |  |  | K-type TC | Bottom | Solex HXer? | F | [5 empty] | 8 | [6 empty] | NI 9213
|  |  |  | K-type TC | Bottom | Solex HXer? | F | [5 empty] | 8 | [6 empty] | NI 9213
|  |  |  | K-type TC | Bottom | Solex HXer? | F | [5 empty] | 8 | [6 empty] | NI 9213
|  |  |  | K-type TC | Bottom | Solex HXer? | F | [5 empty] | 8 | [6 empty] | NI 9213
|  |  |  | K-type TC | Bottom | Solex HXer? | F | [5 empty] | 8 | [6 empty] | NI 9213

## Signals v6
Group | Field Component Vender and Part Number | 2-wire M12 | 4-wire M12 | Field Signal | Location | Version 5 - 2019-05-22 | Units | Eng Min | Eng Max | Raw Min | Raw Max | Wiring Diagram | Mod # | Sig # | Red | Blue | Brown | White | Black | Module | Comments
Relay Outputs |  | 1 |  | 24 VDC | LHS | Pump Enable | Boolean |  |  |  |  | Magnepump v4 | 1 | 0 | [5 empty] | NI 9474
Relay Outputs |  | 2 |  | 24 VDC | LHS | Drive Air Enable | Boolean |  |  |  |  | MAXPRO Filling Cart v4 | 1 | 1 | [5 empty] | NI 9474
Relay Outputs |  | 3 |  | 24 VDC | LHS | Open Vent Valve | Boolean |  |  |  |  | 24 VDC Swagelok Valve v2 | 1 | 2 | [5 empty] | NI 9474 | Convert to single relay?
Relay Outputs |  | 4 |  | 24 VDC | LHS | Open Pressure Trim Valve | Boolean |  |  |  |  | 24 VDC Swagelok Valve v2 | 1 | 3 | [5 empty] | NI 9474 | Convert to single relay?
Relay Outputs |  |  |  | 24 VDC |  | Status Alarm | Boolean | [5 empty] | 1 | 4 | [5 empty] | NI 9474
Relay Outputs |  |  |  | 24 VDC |  | Temperature Alert | Boolean | [5 empty] | 1 | 5 | [5 empty] | NI 9474
Relay Outputs |  |  |  | 24 VDC |  | Open Fill Valve | Boolean |  |  |  |  | 24 VDC Swagelok Valve v2 | 1 | 6 | [5 empty] | NI 9474 | Convert to single relay?
Relay Outputs |  |  |  | 24 VDC |  | 1.5 hp Pump Enable | Boolean |  |  |  |  | 1.5 hp Magnepump v1 | 1 | 7 |  | DO7 |  |  | COM | NI 9474
Cold RTDs | Omega PR-22-3-100-A-1/8-0600-M12 |  |  | 4-wire RTD | RHS | Recup Hot Outlet Temperature [F] | F |  |  |  |  | RTD to NI 9216 | 2 | 0 |  | 2 | 1 | 20 | 21 | NI 9216
Hot RTDs | Conax C16791-1, RTD44W4-INC12-T4(HT)-8.00" |  |  | 4-wire RTD | RHS | Recup Cold Inlet Temperature [F] | F |  |  |  |  | RTD to NI 9216 | 2 | 1 |  | 4 | 3 | 22 | 23 | NI 9216
Hot RTDs | Conax C16791-1, RTD44W4-INC12-T4(HT)-8.00" |  |  | 4-wire RTD | RHS | Heater Outlet Temperature [F] | F |  |  |  |  | RTD to NI 9216 | 2 | 2 |  | 6 | 5 | 24 | 25 | NI 9216
HXer RTDs | Conax C8334-2, RTD43W8-316SS25-T4-8.00" |  |  | 4-wire RTD | RHS | HXer Inlet Temperature [F] | F |  |  |  |  | RTD to NI 9216 | 2 | 3 |  | 8 | 7 | 26 | 27 | NI 9216
HXer RTDs | Conax C8334-2, RTD43W8-316SS25-T4-8.00" |  |  | 4-wire RTD | RHS | HXer Bank 3 West Outlet Temp (B) [F] | F |  |  |  |  | RTD to NI 9216 | 2 | 4 |  | 12 | 11 | 30 | 31 | NI 9216
HXer RTDs | Conax C8334-2, RTD43W8-316SS25-T4-8.00" |  |  | 4-wire RTD | RHS | HXer Bank 3 East Outlet Temp (E) [F] | F |  |  |  |  | RTD to NI 9216 | 2 | 5 |  | 14 | 13 | 32 | 33 | NI 9216
HXer RTDs | Conax C8334-2, RTD43W8-316SS25-T4-8.00" |  |  | 4-wire RTD | RHS | HXer Bank 4 West Outlet Temp (C) [F] | F |  |  |  |  | RTD to NI 9216 | 2 | 6 |  | 16 | 15 | 34 | 35 | NI 9216
HXer RTDs | Conax C8334-2, RTD43W8-316SS25-T4-8.00" |  |  | 4-wire RTD | RHS | HXer Bank 4 East Outlet Temp (F) [F] | F |  |  |  |  | RTD to NI 9216 | 2 | 7 |  | 18 | 17 | 36 | 37 | NI 9216
HXer RTDs | Conax C8334-2, RTD43W8-316SS25-T4-8.00" |  |  | 4-wire RTD | RHS | HXer Bank 2 West Outlet Temp (A) [F] | F |  |  |  |  | RTD to NI 9216 | [7 empty] | NI 9216
HXer RTDs | Conax C8334-2, RTD43W8-316SS25-T4-8.00" |  |  | 4-wire RTD | RHS | HXer Bank 2 East Outlet Temp (D) [F] | F |  |  |  |  | RTD to NI 9216 | [7 empty] | NI 9216
| Conax C8334-2, RTD43W8-316SS25-T4-8.00" |  |  | 4-wire RTD | RHS |  | F |  |  |  |  | RTD to NI 9216 | [7 empty] | NI 9216
|  |  |  | 4-wire RTD | RHS |  | F |  |  |  |  | RTD to NI 9216 | [7 empty] | NI 9216
Cold RTDs | Conax C8334-2, RTD43W8-316SS25-T4-8.00" |  |  | 4-wire RTD | RHS | Cooler Outlet Temperature [F] | F |  |  |  |  | RTD to NI 9219 | 3 | 0 |  | 5 | 3 | 4 | 6 | NI 9219 | Change to NI 9216?
Cold RTDs | Omega PR-22-3-100-A-1/8-0600-M12 |  |  | 4-wire RTD | RHS | Cooler Inlet Temperature [F] | F |  |  |  |  | RTD to NI 9219 | 3 | 1 |  | 5 | 3 | 4 | 6 | NI 9219 | Change to NI 9216?
Water RTDs | Omega PR-22-3-100-A-1/8-0600-M12 |  |  | 4-wire RTD | RHS | Water Inlet Temperature [F] | F |  |  |  |  | RTD to NI 9219 | 3 | 2 |  | 5 | 3 | 4 | 6 | NI 9219 | Change to NI 9216?
Water RTDs | Omega PR-22-3-100-A-1/8-0600-M12 |  |  | 4-wire RTD | RHS | Water Outlet Temperature [F] | F |  |  |  |  | RTD to NI 9219 | 3 | 3 |  | 5 | 3 | 4 | 6 | NI 9219 | Change to NI 9216?
Cold RTDs | WIKA |  |  |  | LHS | 1.5 hp Pump Outlet Temperature [F] | F |  |  |  |  | RTD to NI 9219 | 4 | 0 |  | 5 | 3 | 4 | 6 | NI 9219
Hot RTDs | Conax C17049-1, RTD44W4-INC25-T4(HT)-8.00" |  | 14 | Tri-Loop F #2 | RHS | HXer Outlet Temperature [F] | F |  |  |  |  | RTD to NI 9219 | 4 | 1 |  | 5 | 3 | 4 | 6 | NI 9219
[6 empty] | 1.5 hp Pump dP [psid] | psid | 0 | 27 | 0.004 | 0.02 | Kobold Pressure v3 | 4 | 2 |  |  | Vsup | 3 |  | NI 9219
[6 empty] | 1.5 hp Pump Frequency [Hz] | Hz | 0 | 76.5 | 0.004 | 0.02 | 1.5 hp Magnepump v1 | 4 | 3 |  |  | 3 | 5 |  | NI 9219
mA Outputs |  |  | 1 | 4-20 mA | LHS | Pump Frequency Cmd [Hz] | Hz | 0 | 76.5 | 0.004 | 0.02 | Magnepump v4 | 5 | 0 |  | AO0 |  |  | COM0 | NI 9266
mA Outputs |  |  | 2 | 4-20 mA | LHS | Heater Power Cmd [%] | % | 0 | 100 | 0.004 | 0.02 | KAPL Heater Controller v2 | 5 | 1 |  | AO1 |  |  | COM1 | NI 9266
mA Outputs |  |  | 3 | 4-20 mA | LHS | Recirc Valve % Closed Cmd | % | 0 | 100 | 0.004 | 0.01488 | TRIAC Valve v3 | 5 | 2 |  | AO2 |  |  | COM2 | NI 9266
mA Outputs |  |  | 4 | 4-20 mA | LHS | Flow Valve % Closed Cmd | % | 0 | 100 | 0.004 | 0.02 | TRIAC Valve v3 | 5 | 3 |  | AO3 |  |  | COM3 | NI 9266
mA Outputs |  |  | 5 | 4-20 mA | LHS | Recup Bypass Valve % Closed Cmd | % | 0 | 100 | 0.004 | 0.02 | TRIAC Valve v3 | 5 | 4 |  | AO4 |  |  | COM4 | NI 9266
mA Outputs |  |  | 6 | 4-20 mA | LHS | Drive Air Pressure Cmd | psig | 0 | 141.2708 | 0.004 | 0.02 | MAXPRO Filling Cart v4 | 5 | 5 |  | COM5 | Vsup | AO5 |  | NI 9266
mA Outputs |  |  | 7 | 4-20 mA | LHS | Water Valve % Closed Cmd | % | 0 | 100 | 0.004 | 0.02 |  | 5 | 6 | [5 empty] | NI 9266
mA Outputs |  |  | 8 | 4-20 mA | LHS | 1.5 hp Pump Frequency Cmd [Hz] | Hz | 0 | 76.5 | 0.004 | 0.02 | 1.5 hp Magnepump v1 | 5 | 7 |  | DO7 |  |  | COM | NI 9266
Water Flow |  |  | 17 | 4-20 mA | LHS | Water Flow Rate [gpm] | gpm | 0 | 32 | 0.004 | 0.02 | Water Flow Meter v1 | 6 | 0 |  |  |  | 20 | 1 | NI 9208
Pressures | Rosemount |  | 10 | Tri-Loop B BURST | LHS | Cooler CO2 Inlet Pressure [psig] | psig | 0 | 3626 | 0.004 | 0.02 | Rosemount v4 | 6 | 1 | 21 |  |  |  | 2 | NI 9208
Pressure Drops | Rosemount |  | 10 | Tri-Loop B #1 | LHS | Cooler CO2 dP [psid] | psid | 0 | 36.1273 | 0.004 | 0.02 | Rosemount v4 | 6 | 2 | 22 |  |  |  | 3 | NI 9208
Cold RTDs | Omega PR-22-3-100-A-1/8-0600-M12 |  | 10 | Tri-Loop B #2 | LHS | Pump Outlet Temperature [F] | F | -328 | 1562 | 0.004 | 0.02 | Rosemount v4 | 6 | 3 | 23 |  |  |  | 4 | NI 9208
Feedback Signals |  |  | 1 | 4-20 mA | LHS | Pump Frequency [Hz] | Hz | 0 | 76.5 | 0.004 | 0.02 | Magnepump v4 | 6 | 4 |  |  | 5 | 29 |  | NI 9208
Feedback Signals |  |  | 2 | 4-20 mA | LHS | Heater Power [%] | % | 0 | 100 | 0.004 | 0.02 | KAPL Heater Controller v2 | 6 | 5 |  |  | 6 | 29 |  | NI 9208
Pressures | Rosemount 3051SMV3P14A4R3A11A1JL2Q4GMA1039 |  | 9 | Tri-Loop A BURST | LHS | Pump Outlet Pressure [psig] | psig | 0 | 3626 | 0.004 | 0.02 | Rosemount v4 | 6 | 6 | 26 |  |  |  | 7 | NI 9208
Pressure Drops | Rosemount 3051SMV3P14A4R3A11A1JL2Q4GMA1039 |  | 9 | Tri-Loop A #1 | LHS | Pump dP [psid] | psid | 0 | 150 | 0.004 | 0.02 | Rosemount v4 | 6 | 7 | 27 |  |  |  | 8 | NI 9208
Cold RTDs | Conax C8334-2, RTD43W8-316SS25-T4-8.00" |  | 9 | Tri-Loop A #2 | LHS | Pump Inlet Temperature [F] | F | -328 | 1562 | 0.004 | 0.02 | Rosemount v4 | 6 | 8 | 30 |  |  |  | 11 | NI 9208
Feedback Signals |  |  | 3 | 4-20 mA | LHS | Recirc Valve % Closed | % | 0 | 100 | 0.004 | 0.02 | TRIAC Valve v3 | 6 | 9 |  |  | 12 | 9 |  | NI 9208
Feedback Signals |  |  | 4 | 4-20 mA | LHS | Flow Valve % Closed | % | 0 | 100 | 0.004 | 0.02 | TRIAC Valve v3 | 6 | 10 |  |  | 13 | 9 |  | NI 9208
Feedback Signals |  |  | 5 | 4-20 mA | LHS | Recup Bypass Valve % Closed | % | 0 | 100 | 0.004 | 0.02 | TRIAC Valve v3 | 6 | 11 |  |  | 14 | 9 |  | NI 9208
Flows | Emerson CMFS100P999N2BAEKZX_22757 |  | 16 | 4-20 mA | LHS | Pump Flow Rate [kg¦s] | kg/s | 0 | 6.5277 | 0.004 | 0.02 | Coriolis Flowmeter v3 | 6 | 12 |  | 15 |  |  | 28 | NI 9208
Densities | Emerson CMFS100P999N2BAEKZX_22757 |  | 16 | 4-20 mA | LHS | Pump Outlet Density [kg¦m3] | kg/m^3 | 0 | 5000 | 0.004 | 0.02 | Coriolis Flowmeter v3 | 6 | 13 |  |  | 16 | 28 |  | NI 9208
Water Pressures |  |  | 18 | 4-20 mA | LHS | Water Inlet Pressure [psig] | psig | 0 | 140 | 0.004 | 0.02 | Water Pressure Transducer v1 | 6 | 14 |  | 10 | 36 | 17 |  | NI 9208
Water Pressures |  |  | 19 | 4-20 mA | LHS | Water Outlet Pressure [psig] | psig | 0 | 140 | 0.004 | 0.02 | Water Pressure Transducer v1 | 6 | 15 |  | 10 | 37 | 18 |  | NI 9208
Pressure Drops | Rosemount 3051SMV3P14A4R3A11A1JL2Q4GMA1039 |  | 11 | Tri-Loop C #1 | LHS | Recup Cold dP [psid] | psid | 0 | 150 | 0.004 | 0.02 | Rosemount v4 | 7 | 0 | 20 |  |  |  | 1 | NI 9208
Cold RTDs | Omega PR-22-3-100-A-1/8-0600-M12 |  | 11 | Tri-Loop C #2 | LHS | Recup Cold Outlet Temperature [F] | F | -328 | 1562 | 0.004 | 0.02 | Rosemount v4 | 7 | 1 | 21 |  |  |  | 2 | NI 9208
Pressures | Rosemount 3051SMV3P14A4R3A11A1JL2Q4GMA1039 |  | 11 | Tri-Loop C BURST | LHS | Recup Cold Inlet Pressure [psig] | psig | 0 | 3626 | 0.004 | 0.02 | Rosemount v4 | 7 | 2 | 22 |  |  |  | 3 | NI 9208
Pressure Drops | Rosemount 3051SMV3P14A4R3A11A1JL2Q4GMA1039 |  | 12 | Tri-Loop D #1 | RHS | Recup Hot dP [psid] | psid | 0 | 150 | 0.004 | 0.02 | Rosemount v4 | 7 | 3 | 23 |  |  |  | 4 | NI 9208
Hot RTDs | Conax C16791-1, RTD44W4-INC12-T4(HT)-8.00" |  | 12 | Tri-Loop D #2 | RHS | Recup Hot Inlet Temperature [F] | F | -328 | 1562 | 0.004 | 0.02 | Rosemount v4 | 7 | 4 | 24 |  |  |  | 5 | NI 9208
Pressures | Rosemount 3051SMV3P14A4R3A11A1JL2Q4GMA1039 |  | 12 | Tri-Loop D BURST | RHS | Recup Hot Inlet Pressure [psig] | psig | 0 | 3626 | 0.004 | 0.02 | Rosemount v4 | 7 | 5 | 25 |  |  |  | 6 | NI 9208
Flows | Emerson CMFS100P999N2BAEKZX_22757 |  | 15 | 4-20 mA | LHS | Recup Cold Inlet Flow Rate [kg¦s] | kg/s | 0 | 6.5277 | 0.004 | 0.02 | Coriolis Flowmeter v3 | 7 | 6 |  | 7 |  |  | 9 | NI 9208
Densities | Emerson CMFS100P999N2BAEKZX_22757 |  | 15 | 4-20 mA | LHS | Recup Cold Inlet Density [kg¦m^3] | kg/m^3 | 0 | 5000 | 0.004 | 0.02 | Coriolis Flowmeter v3 | 7 | 7 |  |  | 8 | 9 |  | NI 9208
Pressure Drops | Rosemount 3051SMV3P14A4R3A11A1JL2Q4GMA1039 |  | 14 | Tri-Loop F #1 | RHS | HXer dP [psid] | psid | 0 | 150 | 0.004 | 0.02 | Rosemount v4 | 7 | 8 | 30 |  |  |  | 11 | NI 9208
Pressures | WIKA |  | 22 | 4-20 mA | LHS | Regulated Pressure [psig] | psig | 0 | 3000 | 0.004 | 0.02 | WIKA Pressure v1 | 7 | 9 | 31 |  |  |  | 12 | NI 9208
Pressures | Rosemount 3051SMV3P14A4R3A11A1JL2Q4GMA1039 |  | 14 | Tri-Loop F BURST | RHS | HXer Inlet Pressure [psig] | psi? | 0 | 3626 | 0.004 | 0.02 | Rosemount v4 | 7 | 10 | 32 |  |  |  | 13 | NI 9208
Pressures | WIKA |  | 21 | 4-20 mA | LHS | Supply Pressure [psig] | psig | 0 | 3000 | 0.004 | 0.02 | WIKA Pressure v1 | 7 | 11 | 33 |  |  |  | 14 | NI 9208
Pressure Drops | Rosemount |  | 13 | Tri-Loop E #1 | RHS | Heater dP [psid] | psid | 0 | 36.1273 | 0.004 | 0.02 | Rosemount v4 | 7 | 12 | 34 |  |  |  | 15 | NI 9208
HXer RTDs | Conax C16791-1, RTD44W4-INC12-T4(HT)-8.00" |  | 13 | Tri-Loop E #2 | RHS | Heater Inlet Temperature [F] | F | -328 | 1562 | 0.004 | 0.02 | Rosemount v4 | 7 | 13 | 35 |  |  |  | 16 | NI 9208
Pressures | Rosemount |  | 13 | Tri-Loop E BURST | RHS | Heater Inlet Pressure [psig] | psig | 0 | 3626 | 0.004 | 0.02 | Rosemount v4 | 7 | 14 | 36 |  |  |  | 17 | NI 9208
Feedback Signals |  |  | 7 | 4-20 mA | LHS | Water Valve % Closed | % | 0 | 100 | 0.004 | 0.02 |  | 7 | 15 | [5 empty] | NI 9208
|  |  |  | K-type TC | Bottom | Vent Temperature [F] | F | [12 empty] | NI 9213
|  |  |  | K-type TC | Bottom | Supply Temperature [F] | F | [12 empty] | NI 9213
|  |  |  | K-type TC | Bottom | Regulated Temperature [F] | F | [12 empty] | NI 9213
|  |  |  | K-type TC | Bottom | Compressor Temperature [F] | F | [12 empty] | NI 9213
|  |  |  | K-type TC | Bottom | KAPL Heater Lower RHS [F] | F | [12 empty] | NI 9213
|  |  |  | K-type TC | Bottom | KAPL Heater Lower LHS [F] | F | [12 empty] | NI 9213
|  |  |  | K-type TC | Bottom | KAPL Heater Upper RHS [F] | F | [12 empty] | NI 9213
|  |  |  | K-type TC | Bottom | KAPL Heater Upper LHS [F] | F | [12 empty] | NI 9213
|  |  |  | K-type TC | Bottom | Solex HXer? | F | [12 empty] | NI 9213
|  |  |  | K-type TC | Bottom | Solex HXer? | F | [12 empty] | NI 9213
|  |  |  | K-type TC | Bottom | Solex HXer? | F | [12 empty] | NI 9213
|  |  |  | K-type TC | Bottom | Solex HXer? | F | [12 empty] | NI 9213
|  |  |  | K-type TC | Bottom | Solex HXer? | F | [12 empty] | NI 9213
|  |  |  | K-type TC | Bottom | Solex HXer? | F | [12 empty] | NI 9213
|  |  |  | K-type TC | Bottom | Solex HXer? | F | [12 empty] | NI 9213
|  |  |  | K-type TC | Bottom | Solex HXer? | F | [12 empty] | NI 9213

## Signals v5
Group | Field Component Vender and Part Number | 2-wire M12 | 4-wire M12 | Field Signal | Location | Version 5 - 2019-05-22 | Units | Eng Min | Eng Max | Raw Min | Raw Max | Wiring Diagram | Mod # | Sig # | Red | Blue | Brown | White | Black | Module | Comments
Relay Outputs |  | 1 |  | 24 VDC | LHS | Pump Enable | Boolean |  |  |  |  | Magnepump v4 | 1 | 0 | [5 empty] | NI 9474
Relay Outputs |  | 2 |  | 24 VDC | LHS | Drive Air Enable | Boolean |  |  |  |  | MAXPRO Filling Cart v4 | 1 | 1 | [5 empty] | NI 9474
Relay Outputs |  | 3 |  | 24 VDC | LHS | Open Vent Valve | Boolean |  |  |  |  | 24 VDC Swagelok Valve v1 | 1 | 2 | [5 empty] | NI 9474 | Convert to single relay?
Relay Outputs |  | 4 |  | 24 VDC | LHS | Close Vent Valve | Boolean |  |  |  |  | 24 VDC Swagelok Valve v1 | 1 | 3 | [5 empty] | NI 9474 | Convert to single relay?
Relay Outputs |  |  |  | 24 VDC |  | Status Alarm | Boolean | [12 empty] | NI 9474
Relay Outputs |  |  |  | 24 VDC |  | Temperature Alert | Boolean | [12 empty] | NI 9474
Relay Outputs |  |  |  | 24 VDC |  | Open Fill Valve | Boolean |  |  |  |  | 24 VDC Swagelok Valve v1 | [7 empty] | NI 9474 | Convert to single relay?
Relay Outputs |  |  |  | 24 VDC |  | Close Fill Valve | Boolean |  |  |  |  | 24 VDC Swagelok Valve v1 | [7 empty] | NI 9474 | Convert to single relay?
Cold RTDs | Omega PR-22-3-100-A-1/8-0600-M12 |  |  | 4-wire RTD | RHS | Recup Hot Outlet Temperature [F] | F |  |  |  |  | RTD to NI 9216 | 2 | 0 |  | 2 | 1 | 20 | 21 | NI 9216
Hot RTDs | Conax C16791-1, RTD44W4-INC12-T4(HT)-8.00" |  |  | 4-wire RTD | RHS | Recup Cold Inlet Temperature [F] | F |  |  |  |  | RTD to NI 9216 | 2 | 1 |  | 4 | 3 | 22 | 23 | NI 9216
Hot RTDs | Conax C16791-1, RTD44W4-INC12-T4(HT)-8.00" |  |  | 4-wire RTD | RHS | Heater Outlet Temperature [F] | F |  |  |  |  | RTD to NI 9216 | 2 | 2 |  | 6 | 5 | 24 | 25 | NI 9216
HXer RTDs | Conax C8334-2, RTD43W8-316SS25-T4-8.00" |  |  | 4-wire RTD | RHS | HXer Inlet Temperature [F] | F |  |  |  |  | RTD to NI 9216 | 2 | 3 |  | 8 | 7 | 26 | 27 | NI 9216
HXer RTDs | Conax C8334-2, RTD43W8-316SS25-T4-8.00" |  |  | 4-wire RTD | RHS | HXer Bank 3 West Outlet Temp (B) [F] | F |  |  |  |  | RTD to NI 9216 | 2 | 4 |  | 12 | 11 | 30 | 31 | NI 9216
HXer RTDs | Conax C8334-2, RTD43W8-316SS25-T4-8.00" |  |  | 4-wire RTD | RHS | HXer Bank 3 East Outlet Temp (E) [F] | F |  |  |  |  | RTD to NI 9216 | 2 | 5 |  | 14 | 13 | 32 | 33 | NI 9216
HXer RTDs | Conax C8334-2, RTD43W8-316SS25-T4-8.00" |  |  | 4-wire RTD | RHS | HXer Bank 4 West Outlet Temp (C) [F] | F |  |  |  |  | RTD to NI 9216 | 2 | 6 |  | 16 | 15 | 34 | 35 | NI 9216
HXer RTDs | Conax C8334-2, RTD43W8-316SS25-T4-8.00" |  |  | 4-wire RTD | RHS | HXer Bank 4 East Outlet Temp (F) [F] | F |  |  |  |  | RTD to NI 9216 | 2 | 7 |  | 18 | 17 | 36 | 37 | NI 9216
HXer RTDs | Conax C8334-2, RTD43W8-316SS25-T4-8.00" |  |  | 4-wire RTD | RHS | HXer Bank 2 West Outlet Temp (A) [F] | F |  |  |  |  | RTD to NI 9216 | [7 empty] | NI 9216 | Change to NI 9216
HXer RTDs | Conax C8334-2, RTD43W8-316SS25-T4-8.00" |  |  | 4-wire RTD | RHS | HXer Bank 2 East Outlet Temp (D) [F] | F |  |  |  |  | RTD to NI 9216 | [7 empty] | NI 9216 | Change to NI 9216
| Conax C8334-2, RTD43W8-316SS25-T4-8.00" |  |  | 4-wire RTD | RHS |  | F |  |  |  |  | RTD to NI 9216 | [7 empty] | NI 9216 | Extra?
|  |  |  | 4-wire RTD | RHS |  | F |  |  |  |  | RTD to NI 9216 | [7 empty] | NI 9216 | Extra?
Cold RTDs | Conax C8334-2, RTD43W8-316SS25-T4-8.00" |  |  | 4-wire RTD | RHS | Cooler Outlet Temperature [F] | F |  |  |  |  | RTD to NI 9219 | 3 | 0 |  | 5 | 3 | 4 | 6 | NI 9219 | Change to NI 9216
Cold RTDs | Omega PR-22-3-100-A-1/8-0600-M12 |  |  | 4-wire RTD | RHS | Cooler Inlet Temperature [F] | F |  |  |  |  | RTD to NI 9219 | 3 | 1 |  | 5 | 3 | 4 | 6 | NI 9219 | Change to NI 9216
Water RTDs | Omega PR-22-3-100-A-1/8-0600-M12 |  |  | 4-wire RTD | RHS | Water Inlet Temperature [F] | F |  |  |  |  | RTD to NI 9219 | 3 | 2 |  | 5 | 3 | 4 | 6 | NI 9219 | Change to NI 9216
Water RTDs | Omega PR-22-3-100-A-1/8-0600-M12 |  |  | 4-wire RTD | RHS | Water Outlet Temperature [F] | F |  |  |  |  | RTD to NI 9219 | 3 | 3 |  | 5 | 3 | 4 | 6 | NI 9219 | Change to NI 9216
| WIKA |  |  |  | LHS | [7 empty] | 4 | 0 | [5 empty] | NI 9219
Hot RTDs | Conax C17049-1, RTD44W4-INC25-T4(HT)-8.00" |  | 14 | Tri-Loop F #2 | RHS | HXer Outlet Temperature [F] | F |  |  |  |  | RTD to NI 9219 | 4 | 1 |  | 5 | 3 | 4 | 6 | NI 9219 | Change to NI 9216
[13 empty] | 4 | 2 | [5 empty] | NI 9219
[13 empty] | 4 | 3 | [5 empty] | NI 9219
mA Outputs |  |  | 1 | 4-20 mA | LHS | Pump Frequency Cmd [Hz] | Hz | 0 | 76.5 | 0.004 | 0.02 | Magnepump v4 | 5 | 0 |  | AO0 |  |  | COM0 | NI 9266
mA Outputs |  |  | 2 | 4-20 mA | LHS | Heater Power Cmd [%] | % | 0 | 100 | 0.004 | 0.02 | KAPL Heater Controller v2 | 5 | 1 |  | AO1 |  |  | COM1 | NI 9266
mA Outputs |  |  | 3 | 4-20 mA | LHS | Recirc Valve % Closed Cmd | % | 0 | 100 | 0.004 | 0.01488 | TRIAC Valve v3 | 5 | 2 |  | AO2 |  |  | COM2 | NI 9266
mA Outputs |  |  | 4 | 4-20 mA | LHS | Flow Valve % Closed Cmd | % | 0 | 100 | 0.004 | 0.02 | TRIAC Valve v3 | 5 | 3 |  | AO3 |  |  | COM3 | NI 9266
mA Outputs |  |  | 5 | 4-20 mA | LHS | Recup Bypass Valve % Closed Cmd | % | 0 | 100 | 0.004 | 0.02 | TRIAC Valve v3 | 5 | 4 |  | AO4 |  |  | COM4 | NI 9266
mA Outputs |  |  | 6 | 4-20 mA | LHS | Drive Air Pressure Cmd | psig | 0 | 141.2708 | 0.004 | 0.02 | MAXPRO Filling Cart v4 | 5 | 5 |  | COM5 | Vsup | AO5 |  | NI 9266
mA Outputs |  |  | 7 | 4-20 mA | LHS | Water Valve % Closed Cmd | % | 0 | 100 | 0.004 | 0.02 |  | 5 | 6 | [5 empty] | NI 9266
mA Outputs |  |  | 8 | 4-20 mA | LHS |  |  |  |  | 0.004 | 0.02 | [8 empty] | NI 9266
Water Flow |  |  | 17 | 4-20 mA | LHS | Water Flow Rate [gpm] | gpm | 0 | 32 | 0.004 | 0.02 | Water Flow Meter v1 | 6 | 0 |  |  |  | 20 | 1 | NI 9208
Pressures | Rosemount |  | 10 | Tri-Loop B BURST | LHS | Cooler CO2 Inlet Pressure [psig] | psig | 0 | 3626 | 0.004 | 0.02 | Rosemount v4 | 6 | 1 | 21 |  |  |  | 2 | NI 9208
Pressure Drops | Rosemount |  | 10 | Tri-Loop B #1 | LHS | Cooler CO2 dP [psid] | psid | 0 | 36.1273 | 0.004 | 0.02 | Rosemount v4 | 6 | 2 | 22 |  |  |  | 3 | NI 9208
Cold RTDs | Omega PR-22-3-100-A-1/8-0600-M12 |  | 10 | Tri-Loop B #2 | LHS | Pump Outlet Temperature [F] | F | -328 | 1562 | 0.004 | 0.02 | Rosemount v4 | 6 | 3 | 23 |  |  |  | 4 | NI 9208
Feedback Signals |  |  | 1 | 4-20 mA | LHS | Pump Frequency [Hz] | Hz | 0 | 76.5 | 0.004 | 100 | Magnepump v4 | 6 | 4 |  |  | 5 | 29 |  | NI 9208
Feedback Signals |  |  | 2 | 4-20 mA | LHS | Heater Power [%] | % | 0 | 100 | 0.004 | 0.02 | KAPL Heater Controller v2 | 6 | 5 |  |  | 6 | 29 |  | NI 9208
Pressures | Rosemount 3051SMV3P14A4R3A11A1JL2Q4GMA1039 |  | 9 | Tri-Loop A BURST | LHS | Pump Outlet Pressure [psig] | psig | 0 | 3626 | 0.004 | 0.02 | Rosemount v4 | 6 | 6 | 26 |  |  |  | 7 | NI 9208
Pressure Drops | Rosemount 3051SMV3P14A4R3A11A1JL2Q4GMA1039 |  | 9 | Tri-Loop A #1 | LHS | Pump dP [psid] | psid | 0 | 150 | 0.004 | 0.02 | Rosemount v4 | 6 | 7 | 27 |  |  |  | 8 | NI 9208
Cold RTDs | Conax C8334-2, RTD43W8-316SS25-T4-8.00" |  | 9 | Tri-Loop A #2 | LHS | Pump Inlet Temperature [F] | F | -328 | 1562 | 0.004 | 0.02 | Rosemount v4 | 6 | 8 | 30 |  |  |  | 11 | NI 9208
Feedback Signals |  |  | 3 | 4-20 mA | LHS | Recirc Valve % Closed | % | 0 | 100 | 0.004 | 0.02 | TRIAC Valve v3 | 6 | 9 |  |  | 12 | 9 |  | NI 9208
Feedback Signals |  |  | 4 | 4-20 mA | LHS | Flow Valve % Closed | % | 0 | 100 | 0.004 | 0.02 | TRIAC Valve v3 | 6 | 10 |  |  | 13 | 9 |  | NI 9208
Feedback Signals |  |  | 5 | 4-20 mA | LHS | Recup Bypass Valve % Closed | % | 0 | 100 | 0.004 | 0.02 | TRIAC Valve v3 | 6 | 11 |  |  | 14 | 9 |  | NI 9208
Flows | Emerson CMFS100P999N2BAEKZX_22757 |  | 16 | 4-20 mA | LHS | Pump Flow Rate [kg¦s] | kg/s | 0 | 6.5277 | 0.004 | 0.02 | Coriolis Flowmeter v3 | 6 | 12 |  | 15 |  |  | 28 | NI 9208
Densities | Emerson CMFS100P999N2BAEKZX_22757 |  | 16 | 4-20 mA | LHS | Pump Outlet Density [kg¦m3] | kg/m^3 | 0 | 5000 | 0.004 | 0.02 | Coriolis Flowmeter v3 | 6 | 13 |  |  | 16 | 28 |  | NI 9208
Water Pressures |  |  | 18 | 4-20 mA | LHS | Water Inlet Pressure [psig] | psig | 0 | 140 | 0.004 | 0.02 | Water Pressure Transducer v1 | 6 | 14 |  | 10 | 36 | 17 |  | NI 9208
Water Pressures |  |  | 19 | 4-20 mA | LHS | Water Outlet Pressure [psig] | psig | 0 | 140 | 0.004 | 0.02 | Water Pressure Transducer v1 | 6 | 15 |  | 10 | 37 | 18 |  | NI 9208
Pressure Drops | Rosemount 3051SMV3P14A4R3A11A1JL2Q4GMA1039 |  | 11 | Tri-Loop C #1 | LHS | Recup Cold dP [psid] | psid | 0 | 150 | 0.004 | 0.02 | Rosemount v4 | 7 | 0 | 20 |  |  |  | 1 | NI 9208
Cold RTDs | Omega PR-22-3-100-A-1/8-0600-M12 |  | 11 | Tri-Loop C #2 | LHS | Recup Cold Outlet Temperature [F] | F | -328 | 1562 | 0.004 | 0.02 | Rosemount v4 | 7 | 1 | 21 |  |  |  | 2 | NI 9208
Pressures | Rosemount 3051SMV3P14A4R3A11A1JL2Q4GMA1039 |  | 11 | Tri-Loop C BURST | LHS | Recup Cold Inlet Pressure [psig] | psig | 0 | 3626 | 0.004 | 0.02 | Rosemount v4 | 7 | 2 | 22 |  |  |  | 3 | NI 9208
Pressure Drops | Rosemount 3051SMV3P14A4R3A11A1JL2Q4GMA1039 |  | 12 | Tri-Loop D #1 | RHS | Recup Hot dP [psid] | psid | 0 | 150 | 0.004 | 0.02 | Rosemount v4 | 7 | 3 | 23 |  |  |  | 4 | NI 9208
Hot RTDs | Conax C16791-1, RTD44W4-INC12-T4(HT)-8.00" |  | 12 | Tri-Loop D #2 | RHS | Recup Hot Inlet Temperature [F] | F | -328 | 1562 | 0.004 | 0.02 | Rosemount v4 | 7 | 4 | 24 |  |  |  | 5 | NI 9208
Pressures | Rosemount 3051SMV3P14A4R3A11A1JL2Q4GMA1039 |  | 12 | Tri-Loop D BURST | RHS | Recup Hot Inlet Pressure [psig] | psig | 0 | 3626 | 0.004 | 0.02 | Rosemount v4 | 7 | 5 | 25 |  |  |  | 6 | NI 9208
Flows | Emerson CMFS100P999N2BAEKZX_22757 |  | 15 | 4-20 mA | LHS | Recup Cold Inlet Flow Rate [kg¦s] | kg/s | 0 | 6.5277 | 0.004 | 0.02 | Coriolis Flowmeter v3 | 7 | 6 |  | 7 |  |  | 9 | NI 9208
Densities | Emerson CMFS100P999N2BAEKZX_22757 |  | 15 | 4-20 mA | LHS | Recup Cold Inlet Density [kg¦m^3] | kg/m^3 | 0 | 5000 | 0.004 | 0.02 | Coriolis Flowmeter v3 | 7 | 7 |  |  | 8 | 9 |  | NI 9208
Pressure Drops | Rosemount 3051SMV3P14A4R3A11A1JL2Q4GMA1039 |  | 14 | Tri-Loop F #1 | RHS | HXer dP [psid] | psid | 0 | 150 | 0.004 | 0.02 | Rosemount v4 | 7 | 8 | 30 |  |  |  | 11 | NI 9208
Pressures | WIKA |  | 22 | 4-20 mA | LHS | Regulated Pressure [psig] | psig | 0 | 3000 | 0.004 | 0.02 | WIKA Pressure v1 | 7 | 9 | 31 |  |  |  | 12 | NI 9208
Pressures | Rosemount 3051SMV3P14A4R3A11A1JL2Q4GMA1039 |  | 14 | Tri-Loop F BURST | RHS | HXer Inlet Pressure [psig] | psi? | 0 | 3626 | 0.004 | 0.02 | Rosemount v4 | 7 | 10 | 32 |  |  |  | 13 | NI 9208
Pressures | WIKA |  | 21 | 4-20 mA | LHS | Supply Pressure [psig] | psig | 0 | 3000 | 0.004 | 0.02 | WIKA Pressure v1 | 7 | 11 | 33 |  |  |  | 14 | NI 9208
Pressure Drops | Rosemount |  | 13 | Tri-Loop E #1 | RHS | Heater dP [psid] | psid | 0 | 36.1273 | 0.004 | 0.02 | Rosemount v4 | 7 | 12 | 34 |  |  |  | 15 | NI 9208
HXer RTDs | Conax C16791-1, RTD44W4-INC12-T4(HT)-8.00" |  | 13 | Tri-Loop E #2 | RHS | Heater Inlet Temperature [F] | F | -328 | 1562 | 0.004 | 0.02 | Rosemount v4 | 7 | 13 | 35 |  |  |  | 16 | NI 9208
Pressures | Rosemount |  | 13 | Tri-Loop E BURST | RHS | Heater Inlet Pressure [psig] | psig | 0 | 3626 | 0.004 | 0.02 | Rosemount v4 | 7 | 14 | 36 |  |  |  | 17 | NI 9208
Feedback Signals |  |  | 7 | 4-20 mA | LHS | Water Valve % Closed | % | 0 | 100 | 0.004 | 0.02 |  | 7 | 15 | [5 empty] | NI 9208
|  |  |  | K-type TC | Bottom | Vent Temperature [F] | F | [5 empty] | 8 | [6 empty] | NI 9213
|  |  |  | K-type TC | Bottom | Supply Temperature [F] | F | [5 empty] | 8 | [6 empty] | NI 9213
|  |  |  | K-type TC | Bottom | Regulated Temperature [F] | F | [5 empty] | 8 | [6 empty] | NI 9213
|  |  |  | K-type TC | Bottom | Compressor Temperature [F] | F | [5 empty] | 8 | [6 empty] | NI 9213
|  |  |  | K-type TC | Bottom | KAPL Heater Lower RHS [F] | F | [5 empty] | 8 | [6 empty] | NI 9213
|  |  |  | K-type TC | Bottom | KAPL Heater Lower LHS [F] | F | [5 empty] | 8 | [6 empty] | NI 9213
|  |  |  | K-type TC | Bottom | KAPL Heater Upper RHS [F] | F | [5 empty] | 8 | [6 empty] | NI 9213
|  |  |  | K-type TC | Bottom | KAPL Heater Upper LHS [F] | F | [5 empty] | 8 | [6 empty] | NI 9213
|  |  |  | K-type TC | Bottom | Solex HXer? | F | [5 empty] | 8 | [6 empty] | NI 9213
|  |  |  | K-type TC | Bottom | Solex HXer? | F | [5 empty] | 8 | [6 empty] | NI 9213
|  |  |  | K-type TC | Bottom | Solex HXer? | F | [5 empty] | 8 | [6 empty] | NI 9213
|  |  |  | K-type TC | Bottom | Solex HXer? | F | [5 empty] | 8 | [6 empty] | NI 9213
|  |  |  | K-type TC | Bottom | Solex HXer? | F | [5 empty] | 8 | [6 empty] | NI 9213
|  |  |  | K-type TC | Bottom | Solex HXer? | F | [5 empty] | 8 | [6 empty] | NI 9213
|  |  |  | K-type TC | Bottom | Solex HXer? | F | [5 empty] | 8 | [6 empty] | NI 9213
|  |  |  | K-type TC | Bottom | Solex HXer? | F | [5 empty] | 8 | [6 empty] | NI 9213

## Signals v4
Group | Field Component Vender and Part Number | 2-wire M12 | 4-wire M12 | Field Signal | Location | Version 4 - 2019-05-16 | Units | Eng Min | Eng Max | Raw Min | Raw Max | Wiring Diagram | Mod # | Sig # | Red | Blue | Brown | White | Black | Module | Comments
Relay Outputs |  | 1 |  | 24 VDC | LHS | Pump Enable | Boolean |  |  |  |  | Magnepump v4 | 1 | 0 | [5 empty] | NI 9474
Relay Outputs |  | 2 |  | 24 VDC | LHS | Drive Air Enable | Boolean |  |  |  |  | MAXPRO Filling Cart v4 | 1 | 1 | [5 empty] | NI 9474
Relay Outputs |  | 3 |  | 24 VDC | LHS | Open Vent Valve | Boolean |  |  |  |  | 24 VDC Swagelok Valve v1 | 1 | 2 | [5 empty] | NI 9474 | Convert to single relay?
Relay Outputs |  | 4 |  | 24 VDC | LHS | Close Vent Valve | Boolean |  |  |  |  | 24 VDC Swagelok Valve v1 | 1 | 3 | [5 empty] | NI 9474 | Convert to single relay?
Relay Outputs |  |  |  | 24 VDC |  | Status Alarm | Boolean | [12 empty] | NI 9474
Relay Outputs |  |  |  | 24 VDC |  | Temperature Alert | Boolean | [12 empty] | NI 9474
Relay Outputs |  |  |  | 24 VDC |  | Open Fill Valve | Boolean |  |  |  |  | 24 VDC Swagelok Valve v1 | [7 empty] | NI 9474 | Convert to single relay?
Relay Outputs |  |  |  | 24 VDC |  | Close Fill Valve | Boolean |  |  |  |  | 24 VDC Swagelok Valve v1 | [7 empty] | NI 9474 | Convert to single relay?
Cold RTDs | Omega PR-22-3-100-A-1/8-0600-M12 |  |  | 4-wire RTD | RHS | Recup Hot Outlet Temperature [F] | F |  |  |  |  | RTD to NI 9216 | 2 | 0 |  | 2 | 1 | 20 | 21 | NI 9216
Hot RTDs | Conax C16791-1, RTD44W4-INC12-T4(HT)-8.00" |  |  | 4-wire RTD | RHS | Recup Cold Inlet Temperature [F] | F |  |  |  |  | RTD to NI 9216 | 2 | 1 |  | 4 | 3 | 22 | 23 | NI 9216
Hot RTDs | Conax C16791-1, RTD44W4-INC12-T4(HT)-8.00" |  |  | 4-wire RTD | RHS | Heater Outlet Temperature [F] | F |  |  |  |  | RTD to NI 9216 | 2 | 2 |  | 6 | 5 | 24 | 25 | NI 9216
HXer RTDs | Conax C8334-2, RTD43W8-316SS25-T4-8.00" |  |  | 4-wire RTD | RHS | HXer Outlet Temperature [F] | F |  |  |  |  | RTD to NI 9216 | 2 | 3 |  | 8 | 7 | 26 | 27 | NI 9216 | Confirm if this is inlet or outlet
HXer RTDs | Conax C8334-2, RTD43W8-316SS25-T4-8.00" |  |  | 4-wire RTD | RHS | HXer Bank 3 West Outlet Temp (B) [F] | F |  |  |  |  | RTD to NI 9216 | 2 | 4 |  | 12 | 11 | 30 | 31 | NI 9216
HXer RTDs | Conax C8334-2, RTD43W8-316SS25-T4-8.00" |  |  | 4-wire RTD | RHS | HXer Bank 3 East Outlet Temp (E) [F] | F |  |  |  |  | RTD to NI 9216 | 2 | 5 |  | 14 | 13 | 32 | 33 | NI 9216
HXer RTDs | Conax C8334-2, RTD43W8-316SS25-T4-8.00" |  |  | 4-wire RTD | RHS | HXer Bank 4 West Outlet Temp (C) [F] | F |  |  |  |  | RTD to NI 9216 | 2 | 6 |  | 16 | 15 | 34 | 35 | NI 9216
HXer RTDs | Conax C8334-2, RTD43W8-316SS25-T4-8.00" |  |  | 4-wire RTD | RHS | HXer Bank 4 East Outlet Temp (F) [F] | F |  |  |  |  | RTD to NI 9216 | 2 | 7 |  | 18 | 17 | 36 | 37 | NI 9216
HXer RTDs | Conax C8334-2, RTD43W8-316SS25-T4-8.00" |  |  | 4-wire RTD | RHS | HXer Bank 2 West Outlet Temp (A) [F] | F |  |  |  |  | RTD to NI 9216 | [7 empty] | NI 9216
HXer RTDs | Conax C8334-2, RTD43W8-316SS25-T4-8.00" |  |  | 4-wire RTD | RHS | HXer Bank 2 East Outlet Temp (D) [F] | F |  |  |  |  | RTD to NI 9216 | [7 empty] | NI 9216
| Conax C8334-2, RTD43W8-316SS25-T4-8.00" |  |  | 4-wire RTD | RHS |  | F |  |  |  |  | RTD to NI 9216 | [7 empty] | NI 9216 | Extra?
|  |  |  | 4-wire RTD | RHS |  | F |  |  |  |  | RTD to NI 9216 | [7 empty] | NI 9216 | Extra?
Cold RTDs | Conax C8334-2, RTD43W8-316SS25-T4-8.00" |  |  | 4-wire RTD | RHS | Cooler Outlet Temperature [F] | F |  |  |  |  | RTD to NI 9219 | 3 | 0 |  | 5 | 3 | 4 | 6 | NI 9219 | Change to NI 9216
Cold RTDs | Omega PR-22-3-100-A-1/8-0600-M12 |  |  | 4-wire RTD | RHS | Cooler Inlet Temperature [F] | F |  |  |  |  | RTD to NI 9219 | 3 | 1 |  | 5 | 3 | 4 | 6 | NI 9219 | Change to NI 9216
Water RTDs | Omega PR-22-3-100-A-1/8-0600-M12 |  |  | 4-wire RTD | RHS | Water Inlet Temperature [F] | F |  |  |  |  | RTD to NI 9219 | 3 | 2 |  | 5 | 3 | 4 | 6 | NI 9219 | Change to NI 9216
Water RTDs | Omega PR-22-3-100-A-1/8-0600-M12 |  |  | 4-wire RTD | RHS | Water Outlet Temperature [F] | F |  |  |  |  | RTD to NI 9219 | 3 | 3 |  | 5 | 3 | 4 | 6 | NI 9219 | Change to NI 9216
| WIKA |  |  |  | LHS | [7 empty] | 4 | 0 | [5 empty] | NI 9219
Hot RTDs | Conax C17049-1, RTD44W4-INC25-T4(HT)-8.00" |  | 14 | Tri-Loop F #2 | RHS | HXer Inlet Temperature [F] | F |  |  |  |  | RTD to NI 9219 | 4 | 1 | [5 empty] | NI 9219 | Confirm if this is inlet or outlet
[13 empty] | 4 | 2 | [5 empty] | NI 9219
[13 empty] | 4 | 3 | [5 empty] | NI 9219
mA Outputs |  |  | 1 | 4-20 mA | LHS | Pump Frequency Cmd [Hz] | Hz | 0 | 76.5 | 0.004 | 0.02 | Magnepump v4 | 5 | 0 |  | AO0 |  |  | COM0 | NI 9266
mA Outputs |  |  | 2 | 4-20 mA | LHS | Heater Power Cmd [%] | % | 0 | 100 | 0.004 | 0.02 | KAPL Heater Controller v2 | 5 | 1 |  | AO1 |  |  | COM1 | NI 9266
mA Outputs |  |  | 3 | 4-20 mA | LHS | Recirc Valve % Closed Cmd | % | 0 | 100 | 0.004 | 0.01488 | TRIAC Valve v3 | 5 | 2 |  | AO2 |  |  | COM2 | NI 9266
mA Outputs |  |  | 4 | 4-20 mA | LHS | Flow Valve % Closed Cmd | % | 0 | 100 | 0.004 | 0.02 | TRIAC Valve v3 | 5 | 3 |  | AO3 |  |  | COM3 | NI 9266
mA Outputs |  |  | 5 | 4-20 mA | LHS | Recup Bypass Valve % Closed Cmd | % | 0 | 100 | 0.004 | 0.02 | TRIAC Valve v3 | 5 | 4 |  | AO4 |  |  | COM4 | NI 9266
mA Outputs |  |  | 6 | 4-20 mA | LHS | Drive Air Pressure Cmd | psig | 0 | 141.2708 | 0.004 | 0.02 | MAXPRO Filling Cart v4 | 5 | 5 |  | COM5 | Vsup | AO5 |  | NI 9266
mA Outputs |  |  | 7 | 4-20 mA | LHS | Water Valve % Closed Cmd | % | 0 | 100 | 0.004 | 0.02 |  | 5 | 6 | [5 empty] | NI 9266
mA Outputs |  |  | 8 | 4-20 mA | LHS |  |  |  |  | 0.004 | 0.02 | [8 empty] | NI 9266
Water Flow |  |  | 17 | 4-20 mA | LHS | Water Flow Rate [gpm] | gpm | 0 | 32 | 0.004 | 0.02 | Water Flow Meter v1 | 6 | 0 |  |  |  | 20 | 1 | NI 9208
Pressures | Rosemount |  | 10 | Tri-Loop B BURST | LHS | Cooler CO2 Inlet Pressure [psig] | psig | 0 | 3626 | 0.004 | 0.02 | Rosemount v4 | 6 | 1 | 21 |  |  |  | 2 | NI 9208
Pressure Drops | Rosemount |  | 10 | Tri-Loop B #1 | LHS | Cooler CO2 dP [psid] | psid | 0 | 36.1273 | 0.004 | 0.02 | Rosemount v4 | 6 | 2 | 22 |  |  |  | 3 | NI 9208
Cold RTDs | Omega PR-22-3-100-A-1/8-0600-M12 |  | 10 | Tri-Loop B #2 | LHS | Pump Outlet Temperature [F] | F | -328 | 1562 | 0.004 | 0.02 | Rosemount v4 | 6 | 3 | 23 |  |  |  | 4 | NI 9208
Feedback Signals |  |  | 1 | 4-20 mA | LHS | Pump Frequency [Hz] | Hz | 0 | 76.5 | 0.004 | 100 | Magnepump v4 | 6 | 4 |  |  | 5 | 29 |  | NI 9208
Feedback Signals |  |  | 2 | 4-20 mA | LHS | Heater Power [%] | % | 0 | 100 | 0.004 | 0.02 | KAPL Heater Controller v2 | 6 | 5 |  |  | 6 | 29 |  | NI 9208
Pressures | Rosemount 3051SMV3P14A4R3A11A1JL2Q4GMA1039 |  | 9 | Tri-Loop A BURST | LHS | Pump Outlet Pressure [psig] | psig | 0 | 3626 | 0.004 | 0.02 | Rosemount v4 | 6 | 6 | 26 |  |  |  | 7 | NI 9208
Pressure Drops | Rosemount 3051SMV3P14A4R3A11A1JL2Q4GMA1039 |  | 9 | Tri-Loop A #1 | LHS | Pump dP [psid] | psid | 0 | 150 | 0.004 | 0.02 | Rosemount v4 | 6 | 7 | 27 |  |  |  | 8 | NI 9208
Cold RTDs | Conax C8334-2, RTD43W8-316SS25-T4-8.00" |  | 9 | Tri-Loop A #2 | LHS | Pump Inlet Temperature [F] | F | -328 | 1562 | 0.004 | 0.02 | Rosemount v4 | 6 | 8 | 30 |  |  |  | 11 | NI 9208
Feedback Signals |  |  | 3 | 4-20 mA | LHS | Recirc Valve % Closed | % | 0 | 100 | 0.004 | 0.02 | TRIAC Valve v3 | 6 | 9 |  |  | 12 | 9 |  | NI 9208
Feedback Signals |  |  | 4 | 4-20 mA | LHS | Flow Valve % Closed | % | 0 | 100 | 0.004 | 0.02 | TRIAC Valve v3 | 6 | 10 |  |  | 13 | 9 |  | NI 9208
Feedback Signals |  |  | 5 | 4-20 mA | LHS | Recup Bypass Valve % Closed | % | 0 | 100 | 0.004 | 0.02 | TRIAC Valve v3 | 6 | 11 |  |  | 14 | 9 |  | NI 9208
Flows | Emerson CMFS100P999N2BAEKZX_22757 |  | 16 | 4-20 mA | LHS | Pump Flow Rate [kg¦s] | kg/s | 0 | 6.5277 | 0.004 | 0.02 | Coriolis Flowmeter v3 | 6 | 12 |  | 15 |  |  | 28 | NI 9208
Densities | Emerson CMFS100P999N2BAEKZX_22757 |  | 16 | 4-20 mA | LHS | Pump Outlet Density [kg¦m3] | kg/m^3 | 0 | 5000 | 0.004 | 0.02 | Coriolis Flowmeter v3 | 6 | 13 |  |  | 16 | 28 |  | NI 9208
Water Pressures |  |  | 18 | 4-20 mA | LHS | Water Inlet Pressure [psig] | psig | 0 | 140 | 0.004 | 0.02 | Water Pressure Transducer v1 | 6 | 14 |  | 10 | 36 | 17 |  | NI 9208
Water Pressures |  |  | 19 | 4-20 mA | LHS | Water Outlet Pressure [psig] | psig | 0 | 140 | 0.004 | 0.02 | Water Pressure Transducer v1 | 6 | 15 |  | 10 | 37 | 18 |  | NI 9208
Pressure Drops | Rosemount 3051SMV3P14A4R3A11A1JL2Q4GMA1039 |  | 11 | Tri-Loop C #1 | LHS | Recup Cold dP [psid] | psid | 0 | 150 | 0.004 | 0.02 | Rosemount v4 | 7 | 0 | 20 |  |  |  | 1 | NI 9208
Cold RTDs | Omega PR-22-3-100-A-1/8-0600-M12 |  | 11 | Tri-Loop C #2 | LHS | Recup Cold Outlet Temperature [F] | F | -328 | 1562 | 0.004 | 0.02 | Rosemount v4 | 7 | 1 | 21 |  |  |  | 2 | NI 9208
Pressures | Rosemount 3051SMV3P14A4R3A11A1JL2Q4GMA1039 |  | 11 | Tri-Loop C BURST | LHS | Recup Cold Inlet Pressure [psig] | psig | 0 | 3626 | 0.004 | 0.02 | Rosemount v4 | 7 | 2 | 22 |  |  |  | 3 | NI 9208
Pressure Drops | Rosemount 3051SMV3P14A4R3A11A1JL2Q4GMA1039 |  | 12 | Tri-Loop D #1 | RHS | Recup Hot dP [psid] | psid | 0 | 150 | 0.004 | 0.02 | Rosemount v4 | 7 | 3 | 23 |  |  |  | 4 | NI 9208
Hot RTDs | Conax C16791-1, RTD44W4-INC12-T4(HT)-8.00" |  | 12 | Tri-Loop D #2 | RHS | Recup Hot Inlet Temperature [F] | F | -328 | 1562 | 0.004 | 0.02 | Rosemount v4 | 7 | 4 | 24 |  |  |  | 5 | NI 9208
Pressures | Rosemount 3051SMV3P14A4R3A11A1JL2Q4GMA1039 |  | 12 | Tri-Loop D BURST | RHS | Recup Hot Inlet Pressure [psig] | psig | 0 | 3626 | 0.004 | 0.02 | Rosemount v4 | 7 | 5 | 25 |  |  |  | 6 | NI 9208
Flows | Emerson CMFS100P999N2BAEKZX_22757 |  | 15 | 4-20 mA | LHS | Recup Cold Inlet Flow Rate [kg¦s] | kg/s | 0 | 6.5277 | 0.004 | 0.02 | Coriolis Flowmeter v3 | 7 | 6 |  | 7 |  |  | 9 | NI 9208
Densities | Emerson CMFS100P999N2BAEKZX_22757 |  | 15 | 4-20 mA | LHS | Recup Cold Inlet Density [kg¦m^3] | kg/m^3 | 0 | 5000 | 0.004 | 0.02 | Coriolis Flowmeter v3 | 7 | 7 |  |  | 8 | 9 |  | NI 9208
Pressure Drops | Rosemount 3051SMV3P14A4R3A11A1JL2Q4GMA1039 |  | 14 | Tri-Loop F #1 | RHS | HXer dP [psid] | psid | 0 | 150 | 0.004 | 0.02 | Rosemount v4 | 7 | 8 | [5 empty] | NI 9208
Pressures | WIKA |  | 22 | 4-20 mA | LHS | Regulated Pressure [psig] | psig | 0 | 3000 | 0.004 | 0.02 | WIKA Pressure v1 | 7 | 9 | 31 |  |  |  | 12 | NI 9208
Pressures | Rosemount 3051SMV3P14A4R3A11A1JL2Q4GMA1039 |  | 14 | Tri-Loop F BURST | RHS | HXer Inlet Pressure [psig] | psi? | 0 | 3626 | 0.004 | 0.02 | Rosemount v4 | 7 | 10 | [5 empty] | NI 9208
Pressures | WIKA |  | 21 | 4-20 mA | LHS | Supply Pressure [psig] | psig | 0 | 3000 | 0.004 | 0.02 | WIKA Pressure v1 | 7 | 11 | 33 |  |  |  | 14 | NI 9208
Pressure Drops | Rosemount |  | 13 | Tri-Loop E #1 | RHS | Heater dP [psid] | psid | 0 | 36.1273 | 0.004 | 0.02 | Rosemount v4 | 7 | 12 | 34 |  |  |  | 15 | NI 9208
HXer RTDs | Conax C16791-1, RTD44W4-INC12-T4(HT)-8.00" |  | 13 | Tri-Loop E #2 | RHS | Heater Inlet Temperature [F] | F | -328 | 1562 | 0.004 | 0.02 | Rosemount v4 | 7 | 13 | 35 |  |  |  | 16 | NI 9208
Pressures | Rosemount |  | 13 | Tri-Loop E BURST | RHS | Heater Inlet Pressure [psig] | psig | 0 | 3626 | 0.004 | 0.02 | Rosemount v4 | 7 | 14 | 36 |  |  |  | 17 | NI 9208
Feedback Signals |  |  | 7 | 4-20 mA | LHS | Water Valve % Closed | % | 0 | 100 | 0.004 | 0.02 |  | 7 | 15 | [5 empty] | NI 9208
|  |  |  | K-type TC | Bottom | Vent Temperature [F] | F | [5 empty] | 8 | [6 empty] | NI 9213
|  |  |  | K-type TC | Bottom | Supply Temperature [F] | F | [5 empty] | 8 | [6 empty] | NI 9213
|  |  |  | K-type TC | Bottom | Regulated Temperature [F] | F | [5 empty] | 8 | [6 empty] | NI 9213
|  |  |  | K-type TC | Bottom | Compressor Temperature [F] | F | [5 empty] | 8 | [6 empty] | NI 9213
|  |  |  | K-type TC | Bottom | KAPL Heater Lower RHS [F] | F | [5 empty] | 8 | [6 empty] | NI 9213
|  |  |  | K-type TC | Bottom | KAPL Heater Lower LHS [F] | F | [5 empty] | 8 | [6 empty] | NI 9213
|  |  |  | K-type TC | Bottom | KAPL Heater Upper RHS [F] | F | [5 empty] | 8 | [6 empty] | NI 9213
|  |  |  | K-type TC | Bottom | KAPL Heater Upper LHS [F] | F | [5 empty] | 8 | [6 empty] | NI 9213
|  |  |  | K-type TC | Bottom | Solex HXer? | F | [5 empty] | 8 | [6 empty] | NI 9213
|  |  |  | K-type TC | Bottom | Solex HXer? | F | [5 empty] | 8 | [6 empty] | NI 9213
|  |  |  | K-type TC | Bottom | Solex HXer? | F | [5 empty] | 8 | [6 empty] | NI 9213
|  |  |  | K-type TC | Bottom | Solex HXer? | F | [5 empty] | 8 | [6 empty] | NI 9213
|  |  |  | K-type TC | Bottom | Solex HXer? | F | [5 empty] | 8 | [6 empty] | NI 9213
|  |  |  | K-type TC | Bottom | Solex HXer? | F | [5 empty] | 8 | [6 empty] | NI 9213
|  |  |  | K-type TC | Bottom | Solex HXer? | F | [5 empty] | 8 | [6 empty] | NI 9213
|  |  |  | K-type TC | Bottom | Solex HXer? | F | [5 empty] | 8 | [6 empty] | NI 9213

## Signals v3
Group | Field Component Vender and Part Number | 2-wire M12 | 4-wire M12 | Field Signal | Location | Label | Units | Eng Min | Eng Max | Raw Min | Raw Max | Wiring Diagram | Mod # | Sig # | Red | Blue | Brown | White | Black | Module | Comments
Relay Outputs |  | 1 |  | 24 VDC | LHS | Pump Enable | Boolean |  |  |  |  | Magnepump v4 | 1 | 0 | [5 empty] | NI 9474
Relay Outputs |  | 2 |  | 24 VDC | LHS | Drive Air Enable | Boolean |  |  |  |  | MAXPRO Filling Cart v4 | 1 | 1 | [5 empty] | NI 9474
Relay Outputs |  | 3 |  | 24 VDC | LHS | Open Vent Valve | Boolean |  |  |  |  | 24 VDC Swagelok Valve v1 | 1 | 2 | [5 empty] | NI 9474 | Convert to single relay?
Relay Outputs |  | 4 |  | 24 VDC | LHS | Close Vent Valve | Boolean |  |  |  |  | 24 VDC Swagelok Valve v1 | 1 | 3 | [5 empty] | NI 9474 | Convert to single relay?
Relay Outputs |  |  |  | 24 VDC |  | Status Alarm | Boolean | [12 empty] | NI 9474
Relay Outputs |  |  |  | 24 VDC |  | Temperature Alert | Boolean | [12 empty] | NI 9474
Relay Outputs |  |  |  | 24 VDC |  | Open Fill Valve | Boolean |  |  |  |  | 24 VDC Swagelok Valve v1 | [7 empty] | NI 9474 | Convert to single relay?
Relay Outputs |  |  |  | 24 VDC |  | Close Fill Valve | Boolean |  |  |  |  | 24 VDC Swagelok Valve v1 | [7 empty] | NI 9474 | Convert to single relay?
Cold RTDs | Omega PR-22-3-100-A-1/8-0600-M12 |  |  | 4-wire RTD | RHS | Recup Hot Outlet Temperature [F] | F |  |  |  |  | RTD to NI 9216 | 2 | 0 |  | 2 | 1 | 20 | 21 | NI 9216
Hot RTDs | Conax C16791-1, RTD44W4-INC12-T4(HT)-8.00" |  |  | 4-wire RTD | RHS | Recup Cold Inlet Temperature [F] | F |  |  |  |  | RTD to NI 9216 | 2 | 1 |  | 4 | 3 | 22 | 23 | NI 9216
Hot RTDs | Conax C16791-1, RTD44W4-INC12-T4(HT)-8.00" |  |  | 4-wire RTD | RHS | Heater Outlet Temperature [F] | F |  |  |  |  | RTD to NI 9216 | 2 | 2 |  | 6 | 5 | 24 | 25 | NI 9216
HXer RTDs | Conax C8334-2, RTD43W8-316SS25-T4-8.00" |  |  | 4-wire RTD | RHS | HXer Outlet Temperature [F] | F |  |  |  |  | RTD to NI 9216 | 2 | 3 |  | 8 | 7 | 26 | 27 | NI 9216
HXer RTDs | Conax C8334-2, RTD43W8-316SS25-T4-8.00" |  |  | 4-wire RTD | RHS | HXer Bank 3 West Outlet Temp (B) [F] | F |  |  |  |  | RTD to NI 9216 | 2 | 4 |  | 12 | 11 | 30 | 31 | NI 9216
HXer RTDs | Conax C8334-2, RTD43W8-316SS25-T4-8.00" |  |  | 4-wire RTD | RHS | HXer Bank 3 East Outlet Temp (E) [F] | F |  |  |  |  | RTD to NI 9216 | 2 | 5 |  | 14 | 13 | 32 | 33 | NI 9216
HXer RTDs | Conax C8334-2, RTD43W8-316SS25-T4-8.00" |  |  | 4-wire RTD | RHS | HXer Bank 4 West Outlet Temp (C) [F] | F |  |  |  |  | RTD to NI 9216 | 2 | 6 |  | 16 | 15 | 34 | 35 | NI 9216
HXer RTDs | Conax C8334-2, RTD43W8-316SS25-T4-8.00" |  |  | 4-wire RTD | RHS | HXer Bank 4 East Outlet Temp (F) [F] | F |  |  |  |  | RTD to NI 9216 | 2 | 7 |  | 18 | 17 | 36 | 37 | NI 9216
HXer RTDs | Conax C8334-2, RTD43W8-316SS25-T4-8.00" |  |  | 4-wire RTD | RHS | HXer Bank 2 West Outlet Temp (A) [F] | F |  |  |  |  | RTD to NI 9216 | [7 empty] | NI 9216
HXer RTDs | Conax C8334-2, RTD43W8-316SS25-T4-8.00" |  |  | 4-wire RTD | RHS | HXer Bank 2 East Outlet Temp (D) [F] | F |  |  |  |  | RTD to NI 9216 | [7 empty] | NI 9216
| Conax C8334-2, RTD43W8-316SS25-T4-8.00" |  |  | 4-wire RTD | RHS |  | F |  |  |  |  | RTD to NI 9216 | [7 empty] | NI 9216 | Extra?
|  |  |  | 4-wire RTD | RHS |  | F |  |  |  |  | RTD to NI 9216 | [7 empty] | NI 9216 | Extra?
Cold RTDs | Conax C8334-2, RTD43W8-316SS25-T4-8.00" |  |  | 4-wire RTD | RHS | Cooler Outlet Temperature [F] | F |  |  |  |  | RTD to NI 9219 | 3 | 0 |  | 5 | 3 | 4 | 6 | NI 9219 | Change to NI 9216
Cold RTDs | Omega PR-22-3-100-A-1/8-0600-M12 |  |  | 4-wire RTD | RHS | Cooler Inlet Temperature [F] | F |  |  |  |  | RTD to NI 9219 | 3 | 1 |  | 5 | 3 | 4 | 6 | NI 9219 | Change to NI 9216
Water RTDs | Omega PR-22-3-100-A-1/8-0600-M12 |  |  | 4-wire RTD | RHS | Water Inlet Temperature [F] | F |  |  |  |  | RTD to NI 9219 | 3 | 2 |  | 5 | 3 | 4 | 6 | NI 9219 | Change to NI 9216
Water RTDs | Omega PR-22-3-100-A-1/8-0600-M12 |  |  | 4-wire RTD | RHS | Water Outlet Temperature [F] | F |  |  |  |  | RTD to NI 9219 | 3 | 3 |  | 5 | 3 | 4 | 6 | NI 9219 | Change to NI 9216
| WIKA |  |  | 4-20 mA | LHS | CH0 |  |  |  | 0.004 | 0.02 |  | 4 | 0 | [5 empty] | NI 9219
Hot RTDs | Conax C17049-1, RTD44W4-INC25-T4(HT)-8.00" |  | 14 | Tri-Loop F #2 | RHS | HXer Inlet Temperature [F] | F | -328 | 1562 | 0.004 | 0.02 | Rosemount v4 | 4 | 1 | [5 empty] | NI 9219
Pressure Drops | Rosemount 3051SMV3P14A4R3A11A1JL2Q4GMA1039 |  | 14 | Tri-Loop F #1 | RHS | HXer dP [psid] | psid | 0 | 150 | 0.004 | 0.02 | Rosemount v4 | 4 | 2 | [5 empty] | NI 9219
Pressures | Rosemount 3051SMV3P14A4R3A11A1JL2Q4GMA1039 |  | 14 | Tri-Loop F BURST | RHS | HXer Inlet Pressure [psig] | psi? | 0 | 3626 | 0.004 | 0.02 | Rosemount v4 | 4 | 3 | [5 empty] | NI 9219
mA Outputs |  |  | 1 | 4-20 mA | LHS | Pump Frequency Cmd [Hz] | Hz | 0 | 76.5 | 0.004 | 0.02 | Magnepump v4 | 5 | 0 |  | AO0 |  |  | COM0 | NI 9266
mA Outputs |  |  | 2 | 4-20 mA | LHS | Heater Power Cmd [%] | % | 0 | 100 | 0.004 | 0.02 | KAPL Heater Controller v2 | 5 | 1 |  | AO1 |  |  | COM1 | NI 9266
mA Outputs |  |  | 3 | 4-20 mA | LHS | Recirc Valve % Closed Cmd | % | 0 | 100 | 0.004 | 0.01488 | TRIAC Valve v3 | 5 | 2 |  | AO2 |  |  | COM2 | NI 9266
mA Outputs |  |  | 4 | 4-20 mA | LHS | Flow Valve % Closed Cmd | % | 0 | 100 | 0.004 | 0.02 | TRIAC Valve v3 | 5 | 3 |  | AO3 |  |  | COM3 | NI 9266
mA Outputs |  |  | 5 | 4-20 mA | LHS | Cooling Valve % Closed Cmd | % | 0 | 100 | 0.004 | 0.02 | TRIAC Valve v3 | 5 | 4 |  | AO4 |  |  | COM4 | NI 9266
mA Outputs |  |  | 6 | 4-20 mA | LHS | Drive Air Pressure Cmd | psig | 0 | 141.2708 | 0.004 | 0.02 | MAXPRO Filling Cart v4 | 5 | 5 |  | COM5 | Vsup | AO5 |  | NI 9266
mA Outputs |  |  | 7 | 4-20 mA | LHS | Water Valve % Closed Cmd | % | 0 | 100 | 0.004 | 0.02 |  | 5 | 6 | [5 empty] | NI 9266 | Extra?
mA Outputs |  |  | 8 | 4-20 mA | LHS | Supply CO2 Pressure Cmd | psig | 0 | 1000 | 0.004 | 0.02 |  | 5 | 7 | [5 empty] | NI 9266 | Extra?
Water Flow |  |  | 17 | 4-20 mA | LHS | Water Flow Rate [gpm] | gpm | 0 | 32 | 0.004 | 0.02 | Water Flow Meter v1 | 6 | 0 |  |  |  | 20 | 1 | NI 9208
Pressures | Rosemount |  | 10 | Tri-Loop B BURST | LHS | Cooler CO2 Inlet Pressure [psig] | psi? | 0 | 3626 | 0.004 | 0.02 | Rosemount v4 | 6 | 1 | 21 |  |  |  | 2 | NI 9208
Pressure Drops | Rosemount |  | 10 | Tri-Loop B #1 | LHS | Cooler CO2 dP [psid] | psid | 0 | 36.1273 | 0.004 | 0.02 | Rosemount v4 | 6 | 2 | 22 |  |  |  | 3 | NI 9208
Cold RTDs | Omega PR-22-3-100-A-1/8-0600-M12 |  | 10 | Tri-Loop B #2 | LHS | Pump Outlet Temperature [F] | F | -328 | 1562 | 0.004 | 0.02 | Rosemount v4 | 6 | 3 | 23 |  |  |  | 4 | NI 9208
Feedback Signals |  |  | 1 | 4-20 mA | LHS | Pump Frequency [Hz] | Hz | 0 | 76.5 | 0.004 | 100 | Magnepump v4 | 6 | 4 |  |  | 5 | 29 |  | NI 9208
Feedback Signals |  |  | 2 | 4-20 mA | LHS | Heater Power [%] | % | 0 | 100 | 0.004 | 0.02 | KAPL Heater Controller v2 | 6 | 5 |  |  | 6 | 29 |  | NI 9208
Pressures | Rosemount 3051SMV3P14A4R3A11A1JL2Q4GMA1039 |  | 9 | Tri-Loop A BURST | LHS | Pump Outlet Pressure [psig] | psi? | 0 | 3626 | 0.004 | 0.02 | Rosemount v4 | 6 | 6 | 26 |  |  |  | 7 | NI 9208
Pressure Drops | Rosemount 3051SMV3P14A4R3A11A1JL2Q4GMA1039 |  | 9 | Tri-Loop A #1 | LHS | Pump dP [psid] | psid | 0 | 150 | 0.004 | 0.02 | Rosemount v4 | 6 | 7 | 27 |  |  |  | 8 | NI 9208
Cold RTDs | Conax C8334-2, RTD43W8-316SS25-T4-8.00" |  | 9 | Tri-Loop A #2 | LHS | Pump Inlet Temperature [F] | F | -328 | 1562 | 0.004 | 0.02 | Rosemount v4 | 6 | 8 | 30 |  |  |  | 11 | NI 9208
Feedback Signals |  |  | 3 | 4-20 mA | LHS | Recirc Valve % Closed | % | 0 | 100 | 0.004 | 0.02 | TRIAC Valve v3 | 6 | 9 |  |  | 12 | 9 |  | NI 9208
Feedback Signals |  |  | 4 | 4-20 mA | LHS | Flow Valve % Closed | % | 0 | 100 | 0.004 | 0.02 | TRIAC Valve v3 | 6 | 10 |  |  | 13 | 9 |  | NI 9208
Feedback Signals |  |  | 5 | 4-20 mA | LHS | Cooling Valve % Closed | % | 0 | 100 | 0.004 | 0.02 | TRIAC Valve v3 | 6 | 11 |  |  | 14 | 9 |  | NI 9208
Flows | Emerson CMFS100P999N2BAEKZX_22757 |  | 16 | 4-20 mA | LHS | Pump Flow Rate [kg¦s] | kg/s | 0 | 6.5277 | 0.004 | 0.02 | Coriolis Flowmeter v3 | 6 | 12 |  | 15 |  |  | 28 | NI 9208
Densities | Emerson CMFS100P999N2BAEKZX_22757 |  | 16 | 4-20 mA | LHS | Pump Outlet Density [kg¦m3] | kg/m^3 | 0 | 5000 | 0.004 | 0.02 | Coriolis Flowmeter v3 | 6 | 13 |  |  | 16 | 28 |  | NI 9208
Water Pressures |  |  | 18 | 4-20 mA | LHS | Water Inlet Pressure [psig] | psig | 0 | 140 | 0.004 | 0.02 | Water Pressure Transducer v1 | 6 | 14 |  | 10 | 36 | 17 |  | NI 9208
Water Pressures |  |  | 19 | 4-20 mA | LHS | Water Outlet Pressure [psig] | psig | 0 | 140 | 0.004 | 0.02 | Water Pressure Transducer v1 | 6 | 15 |  | 10 | 37 | 18 |  | NI 9208
Pressure Drops | Rosemount 3051SMV3P14A4R3A11A1JL2Q4GMA1039 |  | 11 | Tri-Loop C #1 | LHS | Recup Cold dP [psid] | psid | 0 | 150 | 0.004 | 0.02 | Rosemount v4 | 7 | 0 | 20 |  |  |  | 1 | NI 9208
Cold RTDs | Omega PR-22-3-100-A-1/8-0600-M12 |  | 11 | Tri-Loop C #2 | LHS | Recup Cold Outlet Temperature [F] | F | -328 | 1562 | 0.004 | 0.02 | Rosemount v4 | 7 | 1 | 21 |  |  |  | 2 | NI 9208
Pressures | Rosemount 3051SMV3P14A4R3A11A1JL2Q4GMA1039 |  | 11 | Tri-Loop C BURST | LHS | Recup Cold Inlet Pressure [psig] | psi? | 0 | 3626 | 0.004 | 0.02 | Rosemount v4 | 7 | 2 | 22 |  |  |  | 3 | NI 9208
Pressure Drops | Rosemount 3051SMV3P14A4R3A11A1JL2Q4GMA1039 |  | 12 | Tri-Loop D #1 | RHS | Recup Hot dP [psid] | psid | 0 | 150 | 0.004 | 0.02 | Rosemount v4 | 7 | 3 | 23 |  |  |  | 4 | NI 9208
Hot RTDs | Conax C16791-1, RTD44W4-INC12-T4(HT)-8.00" |  | 12 | Tri-Loop D #2 | RHS | Recup Hot Inlet Temperature [F] | F | -328 | 1562 | 0.004 | 0.02 | Rosemount v4 | 7 | 4 | 24 |  |  |  | 5 | NI 9208
Pressures | Rosemount 3051SMV3P14A4R3A11A1JL2Q4GMA1039 |  | 12 | Tri-Loop D BURST | RHS | Recup Hot Inlet Pressure [psig] | psi? | 0 | 3626 | 0.004 | 0.02 | Rosemount v4 | 7 | 5 | 25 |  |  |  | 6 | NI 9208
Flows | Emerson CMFS100P999N2BAEKZX_22757 |  | 15 | 4-20 mA | LHS | Recup Cold Inlet Flow Rate [kg¦s] | kg/s | 0 | 6.5277 | 0.004 | 0.02 | Coriolis Flowmeter v3 | 7 | 6 |  | 7 |  |  | 9 | NI 9208
Densities | Emerson CMFS100P999N2BAEKZX_22757 |  | 15 | 4-20 mA | LHS | Recup Cold Inlet Density [kg¦m^3] | kg/m^3 | 0 | 5000 | 0.004 | 0.02 | Coriolis Flowmeter v3 | 7 | 7 |  |  | 8 | 9 |  | NI 9208
Pressure Drops | Kobold PAD-HEE5S2NS00 |  | 24 | 4-20 mA | LHS | Inlet Filter dP [psid] | psid | 0 | 27 | 0.004 | 0.02 | Kobold Pressure v2 | 7 | 8 | 30 |  |  |  | 11 | NI 9208
Pressures | WIKA |  | 22 | 4-20 mA | LHS | Regulated Pressure [psig] | psig | 0 | 3000 | 0.004 | 0.02 | WIKA Pressure v1 | 7 | 9 | 31 |  |  |  | 12 | NI 9208
Pressure Drops | Kobold PAD-HEE5S2NS00 |  | 23 | 4-20 mA | LHS | Outlet Filter dP [psid] | psid | 0 | 27 | 0.004 | 0.02 | Kobold Pressure v2 | 7 | 10 | 32 |  |  |  | 13 | NI 9208
Pressures | WIKA |  | 21 | 4-20 mA | LHS | Supply Pressure [psig] | psig | 0 | 3000 | 0.004 | 0.02 | WIKA Pressure v1 | 7 | 11 | 33 |  |  |  | 14 | NI 9208
Pressure Drops | Rosemount |  | 13 | Tri-Loop E #1 | RHS | Heater dP [psid] | psid | 0 | 36.1273 | 0.004 | 0.02 | Rosemount v4 | 7 | 12 | 34 |  |  |  | 15 | NI 9208
HXer RTDs | Conax C16791-1, RTD44W4-INC12-T4(HT)-8.00" |  | 13 | Tri-Loop E #2 | RHS | Heater Inlet Temperature [F] | F | -328 | 1562 | 0.004 | 0.02 | Rosemount v4 | 7 | 13 | 35 |  |  |  | 16 | NI 9208
Pressures | Rosemount |  | 13 | Tri-Loop E BURST | RHS | Heater Inlet Pressure [psig] | psi? | 0 | 3626 | 0.004 | 0.02 | Rosemount v4 | 7 | 14 | 36 |  |  |  | 17 | NI 9208
Feedback Signals |  |  | 7 | 4-20 mA | LHS | Water Valve % Closed | % | 0 | 100 | 0.004 | 0.02 |  | 7 | 15 | [5 empty] | NI 9208 | Extra?
|  |  |  | K-type TC | Bottom | Vent Temperature [F] | F | [5 empty] | 8 | [6 empty] | NI 9213
|  |  |  | K-type TC | Bottom | Supply Temperature [F] | F | [5 empty] | 8 | [6 empty] | NI 9213
|  |  |  | K-type TC | Bottom | Regulated Temperature [F] | F | [5 empty] | 8 | [6 empty] | NI 9213
|  |  |  | K-type TC | Bottom | Compressor Temperature [F] | F | [5 empty] | 8 | [6 empty] | NI 9213
|  |  |  | K-type TC | Bottom | KAPL Heater Lower RHS [F] | F | [5 empty] | 8 | [6 empty] | NI 9213
|  |  |  | K-type TC | Bottom | KAPL Heater Lower LHS [F] | F | [5 empty] | 8 | [6 empty] | NI 9213
|  |  |  | K-type TC | Bottom | KAPL Heater Upper RHS [F] | F | [5 empty] | 8 | [6 empty] | NI 9213
|  |  |  | K-type TC | Bottom | KAPL Heater Upper LHS [F] | F | [5 empty] | 8 | [6 empty] | NI 9213
|  |  |  | K-type TC | Bottom | Solex HXer? | F | [5 empty] | 8 | [6 empty] | NI 9213
|  |  |  | K-type TC | Bottom | Solex HXer? | F | [5 empty] | 8 | [6 empty] | NI 9213
|  |  |  | K-type TC | Bottom | Solex HXer? | F | [5 empty] | 8 | [6 empty] | NI 9213
|  |  |  | K-type TC | Bottom | Solex HXer? | F | [5 empty] | 8 | [6 empty] | NI 9213
|  |  |  | K-type TC | Bottom | Solex HXer? | F | [5 empty] | 8 | [6 empty] | NI 9213
|  |  |  | K-type TC | Bottom | Solex HXer? | F | [5 empty] | 8 | [6 empty] | NI 9213
|  |  |  | K-type TC | Bottom | Solex HXer? | F | [5 empty] | 8 | [6 empty] | NI 9213
|  |  |  | K-type TC | Bottom | Solex HXer? | F | [5 empty] | 8 | [6 empty] | NI 9213

## Signals v2
Group | Field Component Vender and Part Number | 2-wire M12 | 4-wire M12 | Field Signal | Location | DIN | Label | Units | Eng Min | Eng Max | Raw Min | Raw Max | Wiring Diagram | Mod # | Sig # | Red | Blue | Brown | White | Black | Module | Comments
Relay Outputs |  | 1 |  | 24 VDC | LHS | 27 | Pump Enable | Boolean |  |  |  |  | Magnepump v4 | 1 | 0 |  | CH0b | CH0a |  |  | NI 9474
Relay Outputs |  | 2 |  | 24 VDC | LHS | 29 | Drive Air Enable | Boolean |  |  |  |  | MAXPRO Filling Cart v4 | 1 | 1 | CH1a | COM | CH1b |  |  | NI 9474
Relay Outputs |  | 3 |  | 24 VDC | LHS | 30 | Open Vent Valve | Boolean |  |  |  |  | 24 VDC Swagelok Valve v1 | 1 | 2 | CH2a |  | CH2b |  | COM | NI 9474 | Convert to single relay?
Relay Outputs |  | 4 |  | 24 VDC | LHS | 28 | Close Vent Valve | Boolean |  |  |  |  | 24 VDC Swagelok Valve v1 | 1 | 3 | CH3a |  |  | CH3b | COM | NI 9474
Relay Outputs |  |  |  | 24 VDC |  |  | Status Alarm | Boolean | [12 empty] | NI 9474
Relay Outputs |  |  |  | 24 VDC |  |  | Temperature Alert | Boolean | [12 empty] | NI 9474
Relay Outputs |  |  |  | 24 VDC |  |  | DO6 | Boolean | [12 empty] | NI 9474
Relay Outputs |  |  |  | 24 VDC |  |  | DO7 | Boolean | [12 empty] | NI 9474
Cold RTDs | Omega PR-22-3-100-A-1/8-0600-M12 |  |  | 4-wire RTD | RHS | 5,6 | Recup Hot Outlet Temperature [F] | F |  |  |  |  | RTD to NI 9219 | 2 | 0 |  | 5 | 3 | 4 | 6 | NI 9219 | Change to NI 9216
Hot RTDs | Conax C16791-1, RTD44W4-INC12-T4(HT)-8.00" |  |  | 4-wire RTD | RHS | 7,8 | Recup Cold Inlet Temperature [F] | F |  |  |  |  | RTD to NI 9219 | 2 | 1 |  | 5 | 3 | 4 | 6 | NI 9219 | Change to NI 9216
Hot RTDs | Conax C16791-1, RTD44W4-INC12-T4(HT)-8.00" |  |  | 4-wire RTD | RHS | 11,12 | Heater Outlet Temperature [F] | F |  |  |  |  | RTD to NI 9219 | 2 | 2 |  | 5 | 3 | 4 | 6 | NI 9219 | Need male M12 adpater; Change to NI 9216
Hot RTDs | Conax C8334-2, RTD43W8-316SS25-T4-8.00" |  |  | 4-wire RTD | RHS | 13,14 | Solex HXer Outlet Temperature [F] | F |  |  |  |  | RTD to NI 9219 | 2 | 3 |  | 5 | 3 | 4 | 6 | NI 9219 | Change to NI 9216
Cold RTDs | Conax C8334-2, RTD43W8-316SS25-T4-8.00" |  |  | 4-wire RTD | RHS | 1,2 | Cooler Outlet Temperature [F] | F |  |  |  |  | RTD to NI 9219 | 3 | 0 |  | 5 | 3 | 4 | 6 | NI 9219 | Change to NI 9216
Cold RTDs | Omega PR-22-3-100-A-1/8-0600-M12 |  |  | 4-wire RTD | RHS | 3,4 | Cooler Inlet Temperature [F] | F |  |  |  |  | RTD to NI 9219 | 3 | 1 |  | 5 | 3 | 4 | 6 | NI 9219 | Change to NI 9216
Water RTDs | Omega PR-22-3-100-A-1/8-0600-M12 |  |  | 4-wire RTD | RHS | 25,26 | Water Inlet Temperature [F] | F |  |  |  |  | RTD to NI 9219 | 3 | 2 |  | 5 | 3 | 4 | 6 | NI 9219 | Change to NI 9216
Water RTDs | Omega PR-22-3-100-A-1/8-0600-M12 |  |  | 4-wire RTD | RHS | 27,28 | Water Outlet Temperature [F] | F |  |  |  |  | RTD to NI 9219 | 3 | 3 |  | 5 | 3 | 4 | 6 | NI 9219 | Change to NI 9216
| WIKA |  |  | 4-20 mA | LHS | 26 | CH0 | [6 empty] | 4 | 0 | [5 empty] | NI 9219 | Unnecessary?
Hot RTDs | Conax C17049-1, RTD44W4-INC25-T4(HT)-8.00" |  | 14 | 4-wire RTD | RHS | Tri-Loop F #2 | HXer Inlet Temperature [F] | F | -328 | 1562 | 0.004 | 0.02 | Rosemount v4 | 4 | 1 | [5 empty] | NI 9219
Pressure Drops | Rosemount 3051SMV3P14A4R3A11A1JL2Q4GMA1039 |  | 14 | N/A | RHS | Tri-Loop F #1 | HXer dP [psid] | psid | 0 | 150 | 0.004 | 0.02 | Rosemount v4 | 4 | 2 | [5 empty] | NI 9219
Pressures | Rosemount 3051SMV3P14A4R3A11A1JL2Q4GMA1039 |  | 14 | N/A | RHS | Tri-Loop F BURST | HXer Inlet Pressure [psig] | psi? | 0 | 3626 | 0.004 | 0.02 | Rosemount v4 | 4 | 3 | [5 empty] | NI 9219
mA Outputs |  |  | 1 | 4-20 mA | LHS | 1 | Pump Frequency Cmd [Hz] | Hz | 0 | 76.5 | 0.004 | 0.02 | Magnepump v4 | 5 | 0 |  | AO0 |  |  | COM0 | NI 9266
mA Outputs |  |  | 2 | 4-20 mA | LHS | 6 | Heater Power Cmd [%] | % | 0 | 100 | 0.004 | 0.02 | KAPL Heater Controller v2 | 5 | 1 |  | AO1 |  |  | COM1 | NI 9266
mA Outputs |  |  | 3 | 4-20 mA | LHS | 2 | Recirc Valve % Closed Cmd | % | 0 | 100 | 0.004 | 0.01488 | TRIAC Valve v3 | 5 | 2 |  | AO2 |  |  | COM2 | NI 9266
mA Outputs |  |  | 4 | 4-20 mA | LHS | 3 | Flow Valve % Closed Cmd | % | 0 | 100 | 0.004 | 0.02 | TRIAC Valve v3 | 5 | 3 |  | AO3 |  |  | COM3 | NI 9266
mA Outputs |  |  | 5 | 4-20 mA | LHS | 4 | Cooling Valve % Closed Cmd | % | 0 | 100 | 0.004 | 0.02 | TRIAC Valve v3 | 5 | 4 |  | AO4 |  |  | COM4 | NI 9266
mA Outputs |  |  | 6 | 4-20 mA | LHS | 7 | Drive Air Pressure Cmd | psig | 0 | 141.2708 | 0.004 | 0.02 | MAXPRO Filling Cart v4 | 5 | 5 |  | COM5 | Vsup | AO5 |  | NI 9266
mA Outputs |  |  | 7 | 4-20 mA | LHS | 5 | Water Valve % Closed Cmd | % | 0 | 100 | 0.004 | 0.02 |  | 5 | 6 | [5 empty] | NI 9266
mA Outputs |  |  | 8 | 4-20 mA | LHS | 8 | Supply CO2 Pressure Cmd | psig | 0 | 1000 | 0.004 | 0.02 |  | 5 | 7 | [5 empty] | NI 9266
Water Flow |  |  | 17 | N/A | LHS | 21 | Water Flow Rate [gpm] | gpm | 0 | 50 | 0.004 | 0.02 | Water Flow Meter v1 | 6 | 0 |  |  |  | 20 | 1 | NI 9208
Pressures | Rosemount |  | 10 | N/A | LHS | Tri-Loop B BURST | Cooler CO2 Inlet Pressure [psig] | psi? | 0 | 3626 | 0.004 | 0.02 | Rosemount v4 | 6 | 1 | 21 |  |  |  | 2 | NI 9208
Pressure Drops | Rosemount |  | 10 | N/A | LHS | Tri-Loop B #1 | Cooler CO2 dP [psid] | psid | 0 | 36.1273 | 0.004 | 0.02 | Rosemount v4 | 6 | 2 | 22 |  |  |  | 3 | NI 9208
Cold RTDs | Omega PR-22-3-100-A-1/8-0600-M12 |  | 10 | 4-wire RTD | LHS | Tri-Loop B #2 | Pump Outlet Temperature [F] | F | -328 | 1562 | 0.004 | 0.02 | Rosemount v4 | 6 | 3 | 23 |  |  |  | 4 | NI 9208
Feedback Signals |  |  | 1 | 4-20 mA | LHS | 9 | Pump Frequency [Hz] | Hz | 0 | 76.5 | 0.004 | 100 | Magnepump v4 | 6 | 4 |  |  | 5 | 29 |  | NI 9208
Feedback Signals |  |  | 2 | 4-20 mA | LHS | 14 | Heater Power [%] | % | 0 | 100 | 0.004 | 0.02 | KAPL Heater Controller v2 | 6 | 5 |  |  | 6 | 29 |  | NI 9208
Pressures | Rosemount 3051SMV3P14A4R3A11A1JL2Q4GMA1039 |  | 9 | N/A | LHS | Tri-Loop A BURST | Pump Outlet Pressure [psig] | psi? | 0 | 3626 | 0.004 | 0.02 | Rosemount v4 | 6 | 6 | 26 |  |  |  | 7 | NI 9208
Pressure Drops | Rosemount 3051SMV3P14A4R3A11A1JL2Q4GMA1039 |  | 9 | N/A | LHS | Tri-Loop A #1 | Pump dP [psid] | psid | 0 | 150 | 0.004 | 0.02 | Rosemount v4 | 6 | 7 | 27 |  |  |  | 8 | NI 9208
Cold RTDs | Conax C8334-2, RTD43W8-316SS25-T4-8.00" |  | 9 | 4-wire RTD | LHS | Tri-Loop A #2 | Pump Inlet Temperature [F] | F | -328 | 1562 | 0.004 | 0.02 | Rosemount v4 | 6 | 8 | 30 |  |  |  | 11 | NI 9208
Feedback Signals |  |  | 3 | 4-20 mA | LHS | 10 | Recirc Valve % Closed | % | 0 | 100 | 0.004 | 0.02 | TRIAC Valve v3 | 6 | 9 |  |  | 12 | 9 |  | NI 9208
Feedback Signals |  |  | 4 | 4-20 mA | LHS | 11 | Flow Valve % Closed | % | 0 | 100 | 0.004 | 0.02 | TRIAC Valve v3 | 6 | 10 |  |  | 13 | 9 |  | NI 9208
Feedback Signals |  |  | 5 | 4-20 mA | LHS | 12 | Cooling Valve % Closed | % | 0 | 100 | 0.004 | 0.02 | TRIAC Valve v3 | 6 | 11 |  |  | 14 | 9 |  | NI 9208
Flows | Emerson CMFS100P999N2BAEKZX_22757 |  | 16 | N/A | LHS | 15 | Pump Flow Rate [kg¦s] | kg/s | 0 | 6.5277 | 0.004 | 0.02 | Coriolis Flowmeter v3 | 6 | 12 |  | 15 |  |  | 28 | NI 9208
Densities | Emerson CMFS100P999N2BAEKZX_22757 |  | 16 | N/A | LHS | 16 | Pump Outlet Density [kg¦m3] | kg/m^3 | 0 | 5000 | 0.004 | 0.02 | Coriolis Flowmeter v3 | 6 | 13 |  |  | 16 | 28 |  | NI 9208
Water Pressures |  |  | 18 | N/A | LHS | 22 | Water Inlet Pressure [psig] | psig | 0 | 140 | 0.004 | 0.02 | Water Pressure Transducer v1 | 6 | 14 |  | 10 | 36 | 17 |  | NI 9208
Water Pressures |  |  | 19 | N/A | LHS | 23 | Water Outlet Pressure [psig] | psig | 0 | 140 | 0.004 | 0.02 | Water Pressure Transducer v1 | 6 | 15 |  | 10 | 37 | 18 |  | NI 9208
Pressure Drops | Rosemount 3051SMV3P14A4R3A11A1JL2Q4GMA1039 |  | 11 | N/A | LHS | Tri-Loop C #1 | Recup Cold dP [psid] | psid | 0 | 150 | 0.004 | 0.02 | Rosemount v4 | 7 | 0 | 20 |  |  |  | 1 | NI 9208
Cold RTDs | Omega PR-22-3-100-A-1/8-0600-M12 |  | 11 | 4-wire RTD | LHS | Tri-Loop C #2 | Recup Cold Outlet Temperature [F] | F | -328 | 1562 | 0.004 | 0.02 | Rosemount v4 | 7 | 1 | 21 |  |  |  | 2 | NI 9208
Pressures | Rosemount 3051SMV3P14A4R3A11A1JL2Q4GMA1039 |  | 11 | N/A | LHS | Tri-Loop C BURST | Recup Cold Inlet Pressure [psig] | psi? | 0 | 3626 | 0.004 | 0.02 | Rosemount v4 | 7 | 2 | 22 |  |  |  | 3 | NI 9208
Pressure Drops | Rosemount 3051SMV3P14A4R3A11A1JL2Q4GMA1039 |  | 12 | N/A | RHS | Tri-Loop D #1 | Recup Hot dP [psid] | psid | 0 | 150 | 0.004 | 0.02 | Rosemount v4 | 7 | 3 | 23 |  |  |  | 4 | NI 9208
Hot RTDs | Conax C16791-1, RTD44W4-INC12-T4(HT)-8.00" |  | 12 | 4-wire RTD | RHS | Tri-Loop D #2 | Recup Hot Inlet Temperature [F] | F | -328 | 1562 | 0.004 | 0.02 | Rosemount v4 | 7 | 4 | 24 |  |  |  | 5 | NI 9208
Pressures | Rosemount 3051SMV3P14A4R3A11A1JL2Q4GMA1039 |  | 12 | N/A | RHS | Tri-Loop D BURST | Recup Hot Inlet Pressure [psig] | psi? | 0 | 3626 | 0.004 | 0.02 | Rosemount v4 | 7 | 5 | 25 |  |  |  | 6 | NI 9208
Flows | Emerson CMFS100P999N2BAEKZX_22757 |  | 15 | N/A | LHS | 17 | Flow Rate [kg¦s] | kg/s | 0 | 6.5277 | 0.004 | 0.02 | Coriolis Flowmeter v3 | 7 | 6 |  | 7 |  |  | 9 | NI 9208
Densities | Emerson CMFS100P999N2BAEKZX_22757 |  | 15 | N/A | LHS | 18 | Recup Cold Inlet Density [kg¦m^3] | kg/m^3 | 0 | 5000 | 0.004 | 0.02 | Coriolis Flowmeter v3 | 7 | 7 |  |  | 8 | 9 |  | NI 9208
Pressure Drops | Kobold PAD-HEE5S2NS00 |  | 24 | N/A | LHS | 19 | Inlet Filter dP [psid] | psid | 0 | 27 | 0.004 | 0.02 | Kobold Pressure v2 | 7 | 8 | 30 |  |  |  | 11 | NI 9208
| WIKA |  | 22 | 4-20 mA | LHS | 25 | Regulated Pressure [psig] | psig | 0 | 3000 | 0.004 | 0.02 | WIKA Pressure v1 | 7 | 9 | 31 |  |  |  | 12 | NI 9208
Pressure Drops | Kobold PAD-HEE5S2NS00 |  | 23 | N/A | LHS | 20 | Outlet Filter dP [psid] | psid | 0 | 27 | 0.004 | 0.02 | Kobold Pressure v2 | 7 | 10 | 32 |  |  |  | 13 | NI 9208
| WIKA |  | 21 | 4-20 mA | LHS | 24 | Supply Pressure [psig] | psig | 0 | 3000 | 0.004 | 0.02 | WIKA Pressure v1 | 7 | 11 | 33 |  |  |  | 14 | NI 9208
Pressure Drops | Rosemount |  | 13 | N/A | RHS | Tri-Loop E #1 | Heater dP [psid] | psid | 0 | 36.1273 | 0.004 | 0.02 | Rosemount v4 | 7 | 12 | 34 |  |  |  | 15 | NI 9208
Hot RTDs | Conax C16791-1, RTD44W4-INC12-T4(HT)-8.00" |  | 13 | 4-wire RTD | RHS | Tri-Loop E #2 | Heater Inlet Temperature [F] | F | -328 | 1562 | 0.004 | 0.02 | Rosemount v4 | 7 | 13 | 35 |  |  |  | 16 | NI 9208
Pressures | Rosemount |  | 13 | N/A | RHS | Tri-Loop E BURST | Heater Inlet Pressure [psig] | psi? | 0 | 3626 | 0.004 | 0.02 | Rosemount v4 | 7 | 14 | 36 |  |  |  | 17 | NI 9208
Feedback Signals |  |  | 7 | 4-20 mA | LHS | 13 | Water Valve % Closed | % | 0 | 100 | 0.004 | 0.02 |  | 7 | 15 | [5 empty] | NI 9208
|  |  |  | K-type TC | Bottom | 1 | Vent Temperature [F] | F | [5 empty] | 8 | [6 empty] | NI 9213
|  |  |  | K-type TC | Bottom | 2 | Supply Temperature [F] | F | [5 empty] | 8 | [6 empty] | NI 9213
|  |  |  | K-type TC | Bottom | 3 | Regulated Temperature [F] | F | [5 empty] | 8 | [6 empty] | NI 9213
|  |  |  | K-type TC | Bottom | 4 | Compressor Temperature [F] | F | [5 empty] | 8 | [6 empty] | NI 9213
|  |  |  | K-type TC | Bottom | 5 | KAPL Heater Lower RHS [F] | F | [5 empty] | 8 | [6 empty] | NI 9213
|  |  |  | K-type TC | Bottom | 6 | KAPL Heater Lower LHS [F] | F | [5 empty] | 8 | [6 empty] | NI 9213
|  |  |  | K-type TC | Bottom | 7 | KAPL Heater Upper RHS [F] | F | [5 empty] | 8 | [6 empty] | NI 9213
|  |  |  | K-type TC | Bottom | 8 | KAPL Heater Upper LHS [F] | F | [5 empty] | 8 | [6 empty] | NI 9213
|  |  |  | K-type TC | Bottom | 9 | Solex HXer? | F | [5 empty] | 8 | [6 empty] | NI 9213
|  |  |  | K-type TC | Bottom | 10 | Solex HXer? | F | [5 empty] | 8 | [6 empty] | NI 9213
|  |  |  | K-type TC | Bottom | 11 | Solex HXer? | F | [5 empty] | 8 | [6 empty] | NI 9213
|  |  |  | K-type TC | Bottom | 12 | Solex HXer? | F | [5 empty] | 8 | [6 empty] | NI 9213
|  |  |  | K-type TC | Bottom | 13 | Solex HXer? | F | [5 empty] | 8 | [6 empty] | NI 9213
|  |  |  | K-type TC | Bottom | 14 | Solex HXer? | F | [5 empty] | 8 | [6 empty] | NI 9213
|  |  |  | K-type TC | Bottom | 15 | Solex HXer? | F | [5 empty] | 8 | [6 empty] | NI 9213
|  |  |  | K-type TC | Bottom | 16 | Solex HXer? | F | [5 empty] | 8 | [6 empty] | NI 9213
Hot RTDs | Conax C8334-2, RTD43W8-316SS25-T4-8.00" |  |  | 4-wire RTD | RHS | 31,32 | Mixing Assembly Outlet Temperature | F |  |  |  |  | RTD to NI 9219 |  |  |  | 5 | 3 | 4 | 6 |  | Change to NI 9216
HXer RTDs | Conax C8334-2, RTD43W8-316SS25-T4-8.00" |  |  | 4-wire RTD | RHS | 9,10 | Solex HXer Bank 4 Outlet Temperature A | F |  |  |  |  | RTD to NI 9219 |  |  |  | 5 | 3 | 4 | 6 |  | Change to NI 9216
HXer RTDs | Conax C8334-2, RTD43W8-316SS25-T4-8.00" |  |  | 4-wire RTD | RHS | 15,16 | Solex HXer Bank 4 Outlet Temperature B | F |  |  |  |  | RTD to NI 9219 |  |  |  | 5 | 3 | 4 | 6 |  | Change to NI 9216
HXer RTDs | Conax C8334-2, RTD43W8-316SS25-T4-8.00" |  |  | 4-wire RTD | RHS | 17,18 | Solex HXer Bank 3 Outlet Temperature A | F |  |  |  |  | RTD to NI 9219 |  |  |  | 5 | 3 | 4 | 6 |  | Change to NI 9216
HXer RTDs | Conax C8334-2, RTD43W8-316SS25-T4-8.00" |  |  | 4-wire RTD | RHS | 19,20 | Solex HXer Bank 3 Outlet Temperature B | F |  |  |  |  | RTD to NI 9219 |  |  |  | 5 | 3 | 4 | 6 |  | Change to NI 9216
HXer RTDs | Conax C8334-2, RTD43W8-316SS25-T4-8.00" |  |  | 4-wire RTD | RHS | 21,22 | Solex HXer Bank 2 Outlet Temperature A | F |  |  |  |  | RTD to NI 9219 |  |  |  | 5 | 3 | 4 | 6 |  | Change to NI 9216
HXer RTDs | Conax C8334-2, RTD43W8-316SS25-T4-8.00" |  |  | 4-wire RTD | RHS | 23,24 | Solex HXer Bank 2 Outlet Temperature B | F |  |  |  |  | RTD to NI 9219 |  |  |  | 5 | 3 | 4 | 6 |  | Change to NI 9216
|  |  |  | 4-wire RTD | RHS | 29,30 |  | F |  |  |  |  | RTD to NI 9219 |  |  |  | 5 | 3 | 4 | 6 |  | Change to NI 9216
|  |  |  | 24 VDC | LHS | 31 |  | Boolean | [13 empty] | Remove?
|  |  |  | 24 VDC | LHS | 32 |  | Boolean | [13 empty] | Remove?

## Signals v1
Group | Field Component Vender and Part Number | 2-wire M12 | 4-wire M12 | Field Signal | Location | DIN | Label | Units | Eng Min | Eng Max | Raw Min | Raw Max | Wiring Diagram | Mod # | Sig # | Red | Blue | Brown | White | Black | Module | Comments
mA Outputs |  |  | 1 | 4-20 mA | LHS | 1 | Pump Frequency Cmd [Hz] | Hz | 0 | 76.5 | 0.004 | 0.02 | Magnepump v4 | 1 | 0 |  | AO0 |  |  | COM0 | NI 9266
mA Outputs |  |  | 2 | 4-20 mA | LHS | 6 | Heater Power Cmd [%] | % | 0 | 100 | 0.004 | 0.02 | KAPL Heater Controller v2 | 1 | 1 |  | AO1 |  |  | COM1 | NI 9266
mA Outputs |  |  | 3 | 4-20 mA | LHS | 2 | Recirc Valve % Closed Cmd | % | 0 | 100 | 0.004 | 0.01488 | TRIAC Valve v3 | 1 | 2 |  | AO2 |  |  | COM2 | NI 9266
mA Outputs |  |  | 4 | 4-20 mA | LHS | 3 | Flow Valve % Closed Cmd | % | 0 | 100 | 0.004 | 0.02 | TRIAC Valve v3 | 1 | 3 |  | AO3 |  |  | COM3 | NI 9266
mA Outputs |  |  | 5 | 4-20 mA | LHS | 4 | Cooling Valve % Closed Cmd | % | 0 | 100 | 0.004 | 0.02 | TRIAC Valve v3 | 1 | 4 |  | AO4 |  |  | COM4 | NI 9266
mA Outputs |  |  | 6 | 4-20 mA | LHS | 7 | Drive Air Pressure Cmd | psig | 0 | 141.2708 | 0.004 | 0.02 | MAXPRO Filling Cart v4 | 1 | 5 |  | COM5 | Vsup | AO5 |  | NI 9266
mA Outputs |  |  | 7 | 4-20 mA | LHS | 5 | Water Valve % Closed Cmd | % | 0 | 100 | 0.004 | 0.02 |  | 1 | 6 | [5 empty] | NI 9266
mA Outputs |  |  | 8 | 4-20 mA | LHS | 8 | Supply CO2 Pressure Cmd | psig | 0 | 1000 | 0.004 | 0.02 |  | 1 | 7 | [5 empty] | NI 9266
Relay Outputs |  | 1 |  | 24 VDC | LHS | 27 | Pump Enable | Boolean |  |  |  |  | Magnepump v4 | 2 | 0 |  | CH0b | CH0a |  |  | NI 9485
|  | 2 |  | 24 VDC | LHS | 29 | Drive Air Enable | Boolean |  |  |  |  | MAXPRO Filling Cart v4 | 2 | 1 | CH1a | COM | CH1b |  |  | NI 9485
|  | 3 |  | 24 VDC | LHS | 30 | Open Vent Valve | Boolean |  |  |  |  | 24 VDC Swagelok Valve v1 | 2 | 2 | CH2a |  | CH2b |  | COM | NI 9485 | Convert to single relay
Relay Outputs |  | 4 |  | 24 VDC | LHS | 28 | Close Vent Valve | Boolean |  |  |  |  | 24 VDC Swagelok Valve v1 | 2 | 3 | CH3a |  |  | CH3b | COM | NI 9485 | Change to 'Critical Error' output
Pressure Drops | Rosemount 3051SMV3P14A4R3A11A1JL2Q4GMA1039 |  | 11 | N/A | LHS | Tri-Loop C #1 | Recup Cold dP [psid] | psid | 0 | 150 | 0.004 | 0.02 | Rosemount v4 | 3 | 0 | 20 |  |  |  | 1 | NI 9208
Cold RTDs | Omega PR-22-3-100-A-1/8-0600-M12 |  | 11 | 4-wire RTD | LHS | Tri-Loop C #2 | Recup Cold Outlet Temperature [F] | F | -328 | 1562 | 0.004 | 0.02 | Rosemount v4 | 3 | 1 | 21 |  |  |  | 2 | NI 9208
Pressures | Rosemount 3051SMV3P14A4R3A11A1JL2Q4GMA1039 |  | 11 | N/A | LHS | Tri-Loop C BURST | Recup Cold Inlet Pressure [psig] | psi? | 0 | 3626 | 0.004 | 0.02 | Rosemount v4 | 3 | 2 | 22 |  |  |  | 3 | NI 9208
Pressure Drops | Rosemount 3051SMV3P14A4R3A11A1JL2Q4GMA1039 |  | 12 | N/A | RHS | Tri-Loop D #1 | Recup Hot dP [psid] | psid | 0 | 150 | 0.004 | 0.02 | Rosemount v4 | 3 | 3 | 23 |  |  |  | 4 | NI 9208
Hot RTDs | Conax C16791-1, RTD44W4-INC12-T4(HT)-8.00" |  | 12 | 4-wire RTD | RHS | Tri-Loop D #2 | Recup Hot Inlet Temperature [F] | F | -328 | 1562 | 0.004 | 0.02 | Rosemount v4 | 3 | 4 | 24 |  |  |  | 5 | NI 9208
Pressures | Rosemount 3051SMV3P14A4R3A11A1JL2Q4GMA1039 |  | 12 | N/A | RHS | Tri-Loop D BURST | Recup Hot Inlet Pressure [psig] | psi? | 0 | 3626 | 0.004 | 0.02 | Rosemount v4 | 3 | 5 | 25 |  |  |  | 6 | NI 9208
Flows | Emerson CMFS100P999N2BAEKZX_22757 |  | 15 | N/A | LHS | 17 | Flow Rate [kg¦s] | kg/s | 0 | 6.5277 | 0.004 | 0.02 | Coriolis Flowmeter v3 | 3 | 6 |  | 7 |  |  | 9 | NI 9208
Densities | Emerson CMFS100P999N2BAEKZX_22757 |  | 15 | N/A | LHS | 18 | Recup Cold Inlet Density [kg¦m^3] | kg/m^3 | 0 | 5000 | 0.004 | 0.02 | Coriolis Flowmeter v3 | 3 | 7 |  |  | 8 | 9 |  | NI 9208
Pressure Drops | Kobold PAD-HEE5S2NS00 |  | 24 | N/A | LHS | 19 | Return Filter dP [psid] | psid | 0 | 27 | 0.004 | 0.02 | Kobold Pressure v2 | 3 | 8 | 30 |  |  |  | 11 | NI 9208
| WIKA |  | 22 | 4-20 mA | LHS | 25 | Regulated Pressure [psig] | psig | 0 | 3000 | 0.004 | 0.02 | WIKA Pressure v1 | 3 | 9 |  | 9 | 31 | 12 |  | NI 9208
Pressure Drops | Kobold PAD-HEE5S2NS00 |  | 23 | N/A | LHS | 20 | Outlet Filter dP [psid] | psid | 0 | 27 | 0.004 | 0.02 | Kobold Pressure v2 | 3 | 10 | 32 |  |  |  | 13 | NI 9208
| WIKA |  | 21 | 4-20 mA | LHS | 24 | Supply Pressure [psig] | psig | 0 | 3000 | 0.004 | 0.02 | WIKA Pressure v1 | 3 | 11 | [5 empty] | NI 9208
Pressure Drops | Rosemount |  | 13 | N/A | RHS | Tri-Loop E #1 | Heater dP [psid] | psid | 0 | 36.1273 | 0.004 | 0.02 | Rosemount v4 | 3 | 12 | 34 |  |  |  | 15 | NI 9208
Hot RTDs | Conax C16791-1, RTD44W4-INC12-T4(HT)-8.00" |  | 13 | 4-wire RTD | RHS | Tri-Loop E #2 | Heater Outlet Temperature [F] | F | -328 | 1562 | 0.004 | 0.02 | Rosemount v4 | 3 | 13 | 35 |  |  |  | 16 | NI 9208
Pressures | Rosemount |  | 13 | N/A | RHS | Tri-Loop E BURST | Heater Inlet Pressure [psig] | psi? | 0 | 3626 | 0.004 | 0.02 | Rosemount v4 | 3 | 14 | 36 |  |  |  | 17 | NI 9208
Feedback Signals |  |  | 7 | 4-20 mA | LHS | 13 | Water Valve % Closed | % | 0 | 100 | 0.004 | 0.02 |  | 3 | 15 | [5 empty] | NI 9208
Water Flow |  |  | 17 | N/A | LHS | 21 | Water Flow Rate [gpm] | gpm | 0 | 50 | 0.004 | 0.02 | Water Flow Meter v1 | 4 | 0 | 20 |  |  |  | 1 | NI 9208
Pressures | Rosemount |  | 10 | N/A | LHS | Tri-Loop B BURST | Cooler CO2 Inlet Pressure [psig] | psi? | 0 | 3626 | 0.004 | 0.02 | Rosemount v4 | 4 | 1 | 21 |  |  |  | 2 | NI 9208
Pressure Drops | Rosemount |  | 10 | N/A | LHS | Tri-Loop B #1 | Cooler CO2 dP [psid] | psid | 0 | 36.1273 | 0.004 | 0.02 | Rosemount v4 | 4 | 2 | 22 |  |  |  | 3 | NI 9208
Cold RTDs | Omega PR-22-3-100-A-1/8-0600-M12 |  | 10 | 4-wire RTD | LHS | Tri-Loop B #2 | Pump Outlet Temperature [F] | F | -328 | 1562 | 0.004 | 0.02 | Rosemount v4 | 4 | 3 | 23 |  |  |  | 4 | NI 9208 | Install!
Feedback Signals |  |  | 1 | 4-20 mA | LHS | 9 | Pump Frequency [Hz] | Hz | 0 | 76.5 | 0.004 | 100 | Magnepump v4 | 4 | 4 |  |  | 5 | 29 |  | NI 9208
Feedback Signals |  |  | 2 | 4-20 mA | LHS | 14 | Heater Power [%] | % | 0 | 100 | 0.004 | 0.02 | KAPL Heater Controller v2 | 4 | 5 |  |  | 6 | 29 |  | NI 9208
Pressures | Rosemount 3051SMV3P14A4R3A11A1JL2Q4GMA1039 |  | 9 | N/A | LHS | Tri-Loop A BURST | Pump Outlet Pressure [psig] | psi? | 0 | 3626 | 0.004 | 0.02 | Rosemount v4 | 4 | 6 | 26 |  |  |  | 7 | NI 9208
Pressure Drops | Rosemount 3051SMV3P14A4R3A11A1JL2Q4GMA1039 |  | 9 | N/A | LHS | Tri-Loop A #1 | Pump dP [psid] | psid | 0 | 150 | 0.004 | 0.02 | Rosemount v4 | 4 | 7 | 27 |  |  |  | 8 | NI 9208
Cold RTDs | Conax C8334-2, RTD43W8-316SS25-T4-8.00" |  | 9 | 4-wire RTD | LHS | Tri-Loop A #2 | Pump Inlet Temperature [F] | F | -328 | 1562 | 0.004 | 0.02 | Rosemount v4 | 4 | 8 | 30 |  |  |  | 11 | NI 9208 | Install!
Feedback Signals |  |  | 3 | 4-20 mA | LHS | 10 | Recirc Valve % Closed | % | 0 | 100 | 0.004 | 0.02 | TRIAC Valve v3 | 4 | 9 |  |  | 12 | 9 |  | NI 9208
Feedback Signals |  |  | 4 | 4-20 mA | LHS | 11 | Flow Valve % Closed | % | 0 | 100 | 0.004 | 0.02 | TRIAC Valve v3 | 4 | 10 |  |  | 13 | 9 |  | NI 9208
Feedback Signals |  |  | 5 | 4-20 mA | LHS | 12 | Cooling Valve % Closed | % | 0 | 100 | 0.004 | 0.02 | TRIAC Valve v3 | 4 | 11 |  |  | 14 | 9 |  | NI 9208
Flows | Emerson CMFS100P999N2BAEKZX_22757 |  | 16 | N/A | LHS | 15 | Pump Flow Rate [kg¦s] | kg/s | 0 | 6.5277 | 0.004 | 0.02 | Coriolis Flowmeter v3 | 4 | 12 |  | 15 |  |  | 28 | NI 9208
Densities | Emerson CMFS100P999N2BAEKZX_22757 |  | 16 | N/A | LHS | 16 | Pump Outlet Density [kg¦m3] | kg/m^3 | 0 | 5000 | 0.004 | 0.02 | Coriolis Flowmeter v3 | 4 | 13 |  |  | 16 | 28 |  | NI 9208
Water Pressures |  |  | 18 | N/A | LHS | 22 | Water Inlet Pressure [psig] | psig | 0 | 140 | 0.004 | 0.02 | Water Pressure Transducer v1 | 4 | 14 |  | 9 | 36 | 17 |  | NI 9208
Water Pressures |  |  | 19 | N/A | LHS | 23 | Water Outlet Pressure [psig] | psig | 0 | 140 | 0.004 | 0.02 | Water Pressure Transducer v1 | 4 | 15 |  | 9 | 37 | 18 |  | NI 9208
Cold RTDs | Conax C8334-2, RTD43W8-316SS25-T4-8.00" |  |  | 4-wire RTD | RHS | 1,2 | Cooler Outlet Temperature [F] | F |  |  |  |  | RTD to NI 9219 | 5 | 0 |  | 5 | 3 | 4 | 6 | NI 9219 | Change to NI 9216
Cold RTDs | Omega PR-22-3-100-A-1/8-0600-M12 |  |  | 4-wire RTD | RHS | 3,4 | Cooler Inlet Temperature [F] | F |  |  |  |  | RTD to NI 9219 | 5 | 1 |  | 5 | 3 | 4 | 6 | NI 9219 | Change to NI 9216
Water RTDs | Omega PR-22-3-100-A-1/8-0600-M12 |  |  | 4-wire RTD | RHS | 25,26 | Water Inlet Temperature [F] | F |  |  |  |  | RTD to NI 9219 | 5 | 2 |  | 5 | 3 | 4 | 6 | NI 9219 | Change to NI 9216
Water RTDs | Omega PR-22-3-100-A-1/8-0600-M12 |  |  | 4-wire RTD | RHS | 27,28 | Water Outlet Temperature [F] | F |  |  |  |  | RTD to NI 9219 | 5 | 3 |  | 5 | 3 | 4 | 6 | NI 9219 | Change to NI 9216
Cold RTDs | Omega PR-22-3-100-A-1/8-0600-M12 |  |  | 4-wire RTD | RHS | 5,6 | Recup Hot Outlet Temperature [F] | F |  |  |  |  | RTD to NI 9219 | 6 | 0 |  | 5 | 3 | 4 | 6 | NI 9219 | Change to NI 9216
Hot RTDs | Conax C16791-1, RTD44W4-INC12-T4(HT)-8.00" |  |  | 4-wire RTD | RHS | 7,8 | Recup Cold Inlet Temperature [F] | F |  |  |  |  | RTD to NI 9219 | 6 | 1 |  | 5 | 3 | 4 | 6 | NI 9219 | Change to NI 9216
Hot RTDs | Conax C16791-1, RTD44W4-INC12-T4(HT)-8.00" |  |  | 4-wire RTD | RHS | 11,12 | Heater Inlet Temperature [F] | F |  |  |  |  | RTD to NI 9219 | 6 | 2 |  | 5 | 3 | 4 | 6 | NI 9219 | Need male M12 adpater; Change to NI 9216
Hot RTDs | Conax C8334-2, RTD43W8-316SS25-T4-8.00" |  |  | 4-wire RTD | RHS | 13,14 | Solex HXer Outlet Temperature [F] | F |  |  |  |  | RTD to NI 9219 | 6 | 3 |  | 5 | 3 | 4 | 6 | NI 9219 | Change to NI 9216
Hot RTDs | Conax C17049-1, RTD44W4-INC25-T4(HT)-8.00" |  | 14 | 4-wire RTD | RHS | Tri-Loop F #2 | HXer Inlet Temperature [F] | F | -328 | 1562 | 0.004 | 0.02 | Rosemount v4 | 7 | [6 empty] | NI 9219
Pressures | Rosemount 3051SMV3P14A4R3A11A1JL2Q4GMA1039 |  | 14 | N/A | RHS | Tri-Loop F BURST | HXer Inlet Pressure [psig] | psi? | 0 | 3626 | 0.004 | 0.02 | Rosemount v4 | 7 | [6 empty] | NI 9219
| WIKA |  |  | 4-20 mA | LHS | 26 | [7 empty] | 7 | [6 empty] | NI 9219 | Unnecessary?
Pressure Drops | Rosemount 3051SMV3P14A4R3A11A1JL2Q4GMA1039 |  | 14 | N/A | RHS | Tri-Loop F #1 | HXer dP [psid] | psid | 0 | 150 | 0.004 | 0.02 | Rosemount v4 | 7 | [6 empty] | NI 9219
|  |  |  | K-type TC | Bottom | 1 | Vent Temperature [F] | F | [5 empty] | 8 | [6 empty] | NI 9213
|  |  |  | K-type TC | Bottom | 2 | Supply Temperature [F] | F | [5 empty] | 8 | [6 empty] | NI 9213
|  |  |  | K-type TC | Bottom | 3 | Regulated Temperature [F] | F | [5 empty] | 8 | [6 empty] | NI 9213
|  |  |  | K-type TC | Bottom | 4 | Compressor Temperature [F] | F | [5 empty] | 8 | [6 empty] | NI 9213
|  |  |  | K-type TC | Bottom | 5 | KAPL Heater Lower RHS [F] | F | [5 empty] | 8 | [6 empty] | NI 9213
|  |  |  | K-type TC | Bottom | 6 | KAPL Heater Lower LHS [F] | F | [5 empty] | 8 | [6 empty] | NI 9213
|  |  |  | K-type TC | Bottom | 7 | KAPL Heater Upper RHS [F] | F | [5 empty] | 8 | [6 empty] | NI 9213
|  |  |  | K-type TC | Bottom | 8 | KAPL Heater Upper LHS [F] | F | [5 empty] | 8 | [6 empty] | NI 9213
|  |  |  | K-type TC | Bottom | 9 | Solex HXer? | F | [5 empty] | 8 | [6 empty] | NI 9213
|  |  |  | K-type TC | Bottom | 10 | Solex HXer? | F | [5 empty] | 8 | [6 empty] | NI 9213
|  |  |  | K-type TC | Bottom | 11 | Solex HXer? | F | [5 empty] | 8 | [6 empty] | NI 9213
|  |  |  | K-type TC | Bottom | 12 | Solex HXer? | F | [5 empty] | 8 | [6 empty] | NI 9213
|  |  |  | K-type TC | Bottom | 13 | Solex HXer? | F | [5 empty] | 8 | [6 empty] | NI 9213
|  |  |  | K-type TC | Bottom | 14 | Solex HXer? | F | [5 empty] | 8 | [6 empty] | NI 9213
|  |  |  | K-type TC | Bottom | 15 | Solex HXer? | F | [5 empty] | 8 | [6 empty] | NI 9213
|  |  |  | K-type TC | Bottom | 16 | Solex HXer? | F | [5 empty] | 8 | [6 empty] | NI 9213
Hot RTDs | Conax C8334-2, RTD43W8-316SS25-T4-8.00" |  |  | 4-wire RTD | RHS | 31,32 | Mixing Assembly Outlet Temperature | F |  |  |  |  | RTD to NI 9219 |  |  |  | 5 | 3 | 4 | 6 | NI 9219 | Change to NI 9216
HXer RTDs | Conax C8334-2, RTD43W8-316SS25-T4-8.00" |  |  | 4-wire RTD | RHS | 9,10 | Solex HXer Bank 4 Outlet Temperature A | F |  |  |  |  | RTD to NI 9219 |  |  |  | 5 | 3 | 4 | 6 | NI 9219 | Change to NI 9216
HXer RTDs | Conax C8334-2, RTD43W8-316SS25-T4-8.00" |  |  | 4-wire RTD | RHS | 15,16 | Solex HXer Bank 4 Outlet Temperature B | F |  |  |  |  | RTD to NI 9219 |  |  |  | 5 | 3 | 4 | 6 | NI 9219 | Change to NI 9216
HXer RTDs | Conax C8334-2, RTD43W8-316SS25-T4-8.00" |  |  | 4-wire RTD | RHS | 17,18 | Solex HXer Bank 3 Outlet Temperature A | F |  |  |  |  | RTD to NI 9219 |  |  |  | 5 | 3 | 4 | 6 | NI 9219 | Change to NI 9216
HXer RTDs | Conax C8334-2, RTD43W8-316SS25-T4-8.00" |  |  | 4-wire RTD | RHS | 19,20 | Solex HXer Bank 3 Outlet Temperature B | F |  |  |  |  | RTD to NI 9219 |  |  |  | 5 | 3 | 4 | 6 | NI 9219 | Change to NI 9216
HXer RTDs | Conax C8334-2, RTD43W8-316SS25-T4-8.00" |  |  | 4-wire RTD | RHS | 21,22 | Solex HXer Bank 2 Outlet Temperature A | F |  |  |  |  | RTD to NI 9219 |  |  |  | 5 | 3 | 4 | 6 | NI 9219 | Change to NI 9216
HXer RTDs | Conax C8334-2, RTD43W8-316SS25-T4-8.00" |  |  | 4-wire RTD | RHS | 23,24 | Solex HXer Bank 2 Outlet Temperature B | F |  |  |  |  | RTD to NI 9219 |  |  |  | 5 | 3 | 4 | 6 | NI 9219 | Change to NI 9216
|  |  |  | 4-wire RTD | RHS | 29,30 |  | F |  |  |  |  | RTD to NI 9219 |  |  |  | 5 | 3 | 4 | 6 | NI 9219 | Change to NI 9216
|  |  |  | 24 VDC | LHS | 31 |  | Boolean | [12 empty] | NI 9485 | Remove?
|  |  |  | 24 VDC | LHS | 32 |  | Boolean | [12 empty] | NI 9485 | Remove?

## Signals OLD
Group | Field Component Vender and Part Number | 2-wire M12 | 4-wire M12 | Field Signal | Location | Label | Column1 | Units | Eng Min | Eng Max | Raw Min | Raw Max | Wiring Diagram | Mod # | Sig # | Red | Blue | Brown | White | Black | Module | Comments
|  | 1 |  | 24 VDC | LHS | Pump Enable |  | Boolean |  |  |  |  | Magnepump v4 | 1 | 0 |  | CH0b | CH0a
|  | 2 |  | 24 VDC | LHS | Drive Air Enable |  | Boolean |  |  |  |  | MAXPRO Filling Cart v4 | 1 | 1 | CH1a | COM | CH1b
|  | 3 |  | 24 VDC | LHS | Open Vent Valve |  | Boolean |  |  |  |  | 24 VDC Swagelok Valve v1 | 1 | 2 | CH2a |  | CH2b |  | COM |  | Convert to single relay
|  | 4 |  | 24 VDC | LHS | Close Vent Valve |  | Boolean |  |  |  |  | 24 VDC Swagelok Valve v1 | 1 | 3 | CH3a |  |  | CH3b | COM |  | Change to 'Critical Error' output
|  |  |  | 24 VDC | LHS | Status Alarm |  | Boolean | [5 empty] | 1 | 4 | [6 empty] | Remove?
|  |  |  | 24 VDC | LHS | Temperature Alert |  | Boolean | [5 empty] | 1 | 5 | [6 empty] | Remove?
|  |  |  | 24 VDC |  |  |  | Boolean | [5 empty] | 1 | 6
|  |  |  | 24 VDC |  |  |  | Boolean | [5 empty] | 1 | 7
Cold RTDs | Omega PR-22-3-100-A-1/8-0600-M12 |  |  | 4-wire RTD | RHS | Recup Hot Outlet Temperature [F] |  | F |  |  |  |  | RTD to NI 9219 | 2 | 0 |  | 5 | 3 | 4 | 6 | NI 9219 | Change to NI 9216
Hot RTDs | Conax C16791-1, RTD44W4-INC12-T4(HT)-8.00" |  |  | 4-wire RTD | RHS | Recup Cold Inlet Temperature [F] |  | F |  |  |  |  | RTD to NI 9219 | 2 | 1 |  | 5 | 3 | 4 | 6 | NI 9219 | Change to NI 9216
Hot RTDs | Conax C16791-1, RTD44W4-INC12-T4(HT)-8.00" |  |  | 4-wire RTD | RHS | Heater Outlet Temperature [F] |  | F |  |  |  |  | RTD to NI 9219 | 2 | 2 |  | 5 | 3 | 4 | 6 | NI 9219 | Need male M12 adpater; Change to NI 9216
Hot RTDs | Conax C8334-2, RTD43W8-316SS25-T4-8.00" |  |  | 4-wire RTD | RHS | Solex HXer Outlet Temperature [F] |  | F |  |  |  |  | RTD to NI 9219 | 2 | 3 |  | 5 | 3 | 4 | 6 | NI 9219 | Change to NI 9216
Cold RTDs | Conax C8334-2, RTD43W8-316SS25-T4-8.00" |  |  | 4-wire RTD | RHS | Cooler Outlet Temperature [F] |  | F |  |  |  |  | RTD to NI 9219 | 3 | 0 |  | 5 | 3 | 4 | 6 | NI 9219 | Change to NI 9216
Cold RTDs | Omega PR-22-3-100-A-1/8-0600-M12 |  |  | 4-wire RTD | RHS | Cooler Inlet Temperature [F] |  | F |  |  |  |  | RTD to NI 9219 | 3 | 1 |  | 5 | 3 | 4 | 6 | NI 9219 | Change to NI 9216
Water RTDs | Omega PR-22-3-100-A-1/8-0600-M12 |  |  | 4-wire RTD | RHS | Water Inlet Temperature [F] |  | F |  |  |  |  | RTD to NI 9219 | 3 | 2 |  | 5 | 3 | 4 | 6 | NI 9219 | Change to NI 9216
Water RTDs | Omega PR-22-3-100-A-1/8-0600-M12 |  |  | 4-wire RTD | RHS | Water Outlet Temperature [F] |  | F |  |  |  |  | RTD to NI 9219 | 3 | 3 |  | 5 | 3 | 4 | 6 | NI 9219 | Change to NI 9216
| WIKA |  |  | 4-20 mA | LHS | [8 empty] | 4 | 0 | [5 empty] | NI 9219 | Unnecessary?
Hot RTDs | Conax C17049-1, RTD44W4-INC25-T4(HT)-8.00" |  | 14 | 4-wire RTD | RHS | HXer Inlet Temperature [F] |  | F | -329.414 | 1546.33 | 0.004 | 0.02 | Rosemount v4 | 4 | 1 |  |  |  |  | 5 | NI 9219 | Tri-Loop F #2
Pressure Drops | Rosemount 3051SMV3P14A4R3A11A1JL2Q4GMA1039 |  | 14 | N/A | RHS | HXer dP [psid] |  | psid | 0 | 150 | 0.004 | 0.02 | Rosemount v4 | 4 | 2 |  |  |  |  | 5 | NI 9219 | Tri-Loop F #1
Pressures | Rosemount 3051SMV3P14A4R3A11A1JL2Q4GMA1039 |  | 14 | N/A | RHS | HXer Inlet Pressure [psig] |  | psi? | 0 | 3626 | 0.004 | 0.02 | Rosemount v4 | 4 | 3 | 3 |  |  |  | 5 | NI 9219 | Tri-Loop F BURST
mA Outputs |  |  | 1 | 4-20 mA | LHS | Pump Frequency Cmd [Hz] |  | Hz | 0 | 76.5 | 0.004 | 0.02 | Magnepump v4 | 5 | 0 |  | AO0 |  |  | COM0 | NI 9266
mA Outputs |  |  | 2 | 4-20 mA | LHS | Heater Power Cmd [%] |  | % | 0 | 100 | 0.004 | 0.02 | KAPL Heater Controller v2 | 5 | 1 |  | AO1 |  |  | COM1 | NI 9266
mA Outputs |  |  | 3 | 4-20 mA | LHS | Recirc Valve % Closed Cmd |  | % | 0 | 100 | 0.004 | 0.01488 | TRIAC Valve v3 | 5 | 2 |  | AO2 |  |  | COM2 | NI 9266
mA Outputs |  |  | 4 | 4-20 mA | LHS | Flow Valve % Closed Cmd |  | % | 0 | 100 | 0.004 | 0.02 | TRIAC Valve v3 | 5 | 3 |  | AO3 |  |  | COM3 | NI 9266
mA Outputs |  |  | 5 | 4-20 mA | LHS | Cooling Valve % Closed Cmd |  | % | 0 | 100 | 0.004 | 0.02 | TRIAC Valve v3 | 5 | 4 |  | AO4 |  |  | COM4 | NI 9266
mA Outputs |  |  | 6 | 4-20 mA | LHS | Drive Air Pressure Cmd |  | psig | 0 | 141.2708 | 0.004 | 0.02 | MAXPRO Filling Cart v4 | 5 | 5 |  | COM5 | Vsup | AO5 |  | NI 9266
mA Outputs |  |  | 7 | 4-20 mA | LHS | Water Valve % Closed Cmd |  | % | 0 | 100 | 0.004 | 0.02 |  | 5 | 6 | [5 empty] | NI 9266
mA Outputs |  |  | 8 | 4-20 mA | LHS | Supply CO2 Pressure Cmd |  | psig | 0 | 1000 | 0.004 | 0.02 |  | 5 | 7 | [5 empty] | NI 9266
Water Flow |  |  | 17 | N/A | LHS | Water Flow Rate [gpm] |  | gpm | 0 | 50 | 0.004 | 0.02 | Water Flow Meter v1 | 6 | 0 |  |  |  | 20 | 1 | NI 9208
Pressures | Rosemount |  | 10 | N/A | LHS | Cooler CO2 Inlet Pressure [psig] |  | psi? | 0 | 3626 | 0.004 | 0.02 | Rosemount v4 | 6 | 1 | 21 |  |  |  | 2 | NI 9208 | Tri-Loop B BURST
Pressure Drops | Rosemount |  | 10 | N/A | LHS | Cooler CO2 dP [psid] |  | psid | 0 | 36.1273 | 0.004 | 0.02 | Rosemount v4 | 6 | 2 | 22 |  |  |  | 3 | NI 9208 | Tri-Loop B #1
Cold RTDs | Omega PR-22-3-100-A-1/8-0600-M12 |  | 10 | 4-wire RTD | LHS | Pump Outlet Temperature [F] |  | F | -332 | 1536.6 | 0.004 | 0.02 | Rosemount v4 | 6 | 3 | 23 |  |  |  | 4 | NI 9208 | Tri-Loop B #2
Feedback Signals |  |  | 1 | 4-20 mA | LHS | Pump Frequency [Hz] |  | Hz | 0 | 76.5 | 0.004 | 0.02 | Magnepump v4 | 6 | 4 |  |  | 5 | 29 |  | NI 9208
Feedback Signals |  |  | 2 | 4-20 mA | LHS | Heater Power [%] |  | % | 0 | 100 | 0.004 | 0.02 | KAPL Heater Controller v2 | 6 | 5 |  |  | 6 | 29 |  | NI 9208
Pressures | Rosemount 3051SMV3P14A4R3A11A1JL2Q4GMA1039 |  | 9 | N/A | LHS | Pump Outlet Pressure [psig] |  | psi? | 0 | 3626 | 0.004 | 0.02 | Rosemount v4 | 6 | 6 | 26 |  |  |  | 7 | NI 9208 | Tri-Loop A BURST
Pressure Drops | Rosemount 3051SMV3P14A4R3A11A1JL2Q4GMA1039 |  | 9 | N/A | LHS | Pump dP [psid] |  | psid | 0 | 150 | 0.004 | 0.02 | Rosemount v4 | 6 | 7 | 27 |  |  |  | 8 | NI 9208 | Tri-Loop A #1
Cold RTDs | Conax C8334-2, RTD43W8-316SS25-T4-8.00" |  | 9 | 4-wire RTD | LHS | Pump Inlet Temperature [F] |  | F | -329.414 | 1546.33 | 0.004 | 0.02 | Rosemount v4 | 6 | 8 | 30 |  |  |  | 11 | NI 9208 | Tri-Loop A #2
Feedback Signals |  |  | 3 | 4-20 mA | LHS | Recirc Valve % Closed |  | % | 0 | 100 | 0.004 | 0.02 | TRIAC Valve v3 | 6 | 9 |  |  | 12 | 9 |  | NI 9208
Feedback Signals |  |  | 4 | 4-20 mA | LHS | Flow Valve % Closed |  | % | 0 | 100 | 0.004 | 0.02 | TRIAC Valve v3 | 6 | 10 |  |  | 13 | 9 |  | NI 9208
Feedback Signals |  |  | 5 | 4-20 mA | LHS | Cooling Valve % Closed |  | % | 0 | 100 | 0.004 | 0.02 | TRIAC Valve v3 | 6 | 11 |  |  | 14 | 9 |  | NI 9208
Flows | Emerson CMFS100P999N2BAEKZX_22757 |  | 16 | N/A | LHS | Pump Flow Rate [kg¦s] |  | kg/s | 0 | 6.5277 | 0.004 | 0.02 | Coriolis Flowmeter v3 | 6 | 12 |  | 15 |  |  | 28 | NI 9208
Densities | Emerson CMFS100P999N2BAEKZX_22757 |  | 16 | N/A | LHS | Pump Outlet Density [kg¦m3] |  | kg/m^3 | 0 | 5000 | 0.004 | 0.02 | Coriolis Flowmeter v3 | 6 | 13 |  |  | 16 | 28 |  | NI 9208
Water Pressures |  |  | 18 | N/A | LHS | Water Inlet Pressure [psig] |  | psig | 0 | 140 | 0.004 | 0.02 | Water Pressure Transducer v1 | 6 | 14 |  | 10 | 36 | 17 |  | NI 9208
Water Pressures |  |  | 19 | N/A | LHS | Water Outlet Pressure [psig] |  | psig | 0 | 140 | 0.004 | 0.02 | Water Pressure Transducer v1 | 6 | 15 |  | 10 | 37 | 18 |  | NI 9208
Pressure Drops | Rosemount 3051SMV3P14A4R3A11A1JL2Q4GMA1039 |  | 11 | N/A | LHS | Recup Cold dP [psid] |  | psid | 0 | 150 | 0.004 | 0.02 | Rosemount v4 | 7 | 0 | 20 |  |  |  | 1 | NI 9208 | Tri-Loop C #1
Cold RTDs | Omega PR-22-3-100-A-1/8-0600-M12 |  | 11 | 4-wire RTD | LHS | Recup Cold Outlet Temperature [F] |  | F | -329.414 | 1546.33 | 0.004 | 0.02 | Rosemount v4 | 7 | 1 | 21 |  |  |  | 2 | NI 9208 | Tri-Loop C #2
Pressures | Rosemount 3051SMV3P14A4R3A11A1JL2Q4GMA1039 |  | 11 | N/A | LHS | Recup Cold Inlet Pressure [psig] |  | psi? | 0 | 3626 | 0.004 | 0.02 | Rosemount v4 | 7 | 2 | 22 |  |  |  | 3 | NI 9208 | Tri-Loop C BURST
Pressure Drops | Rosemount 3051SMV3P14A4R3A11A1JL2Q4GMA1039 |  | 12 | N/A | RHS | Recup Hot dP [psid] |  | psid | 0 | 150 | 0.004 | 0.02 | Rosemount v4 | 7 | 3 | 23 |  |  |  | 4 | NI 9208 | Tri-Loop D #1
Hot RTDs | Conax C16791-1, RTD44W4-INC12-T4(HT)-8.00" |  | 12 | 4-wire RTD | RHS | Recup Hot Inlet Temperature [F] |  | F | -329.414 | 1546.33 | 0.004 | 0.02 | Rosemount v4 | 7 | 4 | 24 |  |  |  | 5 | NI 9208 | Tri-Loop D #2
Pressures | Rosemount 3051SMV3P14A4R3A11A1JL2Q4GMA1039 |  | 12 | N/A | RHS | Recup Hot Inlet Pressure [psig] |  | psi? | 0 | 3626 | 0.004 | 0.02 | Rosemount v4 | 7 | 5 | 25 |  |  |  | 6 | NI 9208 | Tri-Loop D BURST
Flows | Emerson CMFS100P999N2BAEKZX_22757 |  | 15 | N/A | LHS | Flow Rate [kg¦s] |  | kg/s | 0 | 6.5277 | 0.004 | 0.02 | Coriolis Flowmeter v3 | 7 | 6 |  | 7 |  |  | 9 | NI 9208
Densities | Emerson CMFS100P999N2BAEKZX_22757 |  | 15 | N/A | LHS | Recup Cold Inlet Density [kg¦m^3] |  | kg/m^3 | 0 | 5000 | 0.004 | 0.02 | Coriolis Flowmeter v3 | 7 | 7 |  |  | 8 | 9 |  | NI 9208
Pressure Drops | Kobold PAD-HEE5S2NS00 |  | 24 | N/A | LHS | Inlet Filter dP [psid] |  | psid | 0 | 27 | 0.004 | 0.02 | Kobold Pressure v2 | 7 | 8 | 30 |  |  |  | 11 | NI 9208
| WIKA |  | 22 | 4-20 mA | LHS | Regulated Pressure [psig] |  | psig | 0 | 3000 | 0.004 | 0.02 | WIKA Pressure v1 | 7 | 9 | 31 |  |  |  | 12 | NI 9208
Pressure Drops | Kobold PAD-HEE5S2NS00 |  | 23 | N/A | LHS | Outlet Filter dP [psid] |  | psid | 0 | 27 | 0.004 | 0.02 | Kobold Pressure v2 | 7 | 10 | 32 |  |  |  | 13 | NI 9208
| WIKA |  | 21 | 4-20 mA | LHS | Supply Pressure [psig] |  | psig | 0 | 3000 | 0.004 | 0.02 | WIKA Pressure v1 | 7 | 11 | 33 |  |  |  | 14 | NI 9208
Pressure Drops | Rosemount |  | 13 | N/A | RHS | Heater dP [psid] |  | psid | 0 | 36.1273 | 0.004 | 0.02 | Rosemount v4 | 7 | 12 | 34 |  |  |  | 15 | NI 9208 | Tri-Loop E #1
Hot RTDs | Conax C16791-1, RTD44W4-INC12-T4(HT)-8.00" |  | 13 | 4-wire RTD | RHS | Heater Inlet Temperature [F] |  | F | -332 | 1536.6 | 0.004 | 0.02 | Rosemount v4 | 7 | 13 | 35 |  |  |  | 16 | NI 9208 | Tri-Loop E #2
Pressures | Rosemount |  | 13 | N/A | RHS | Heater Inlet Pressure [psig] |  | psi? | 0 | 3626 | 0.004 | 0.02 | Rosemount v4 | 7 | 14 | 36 |  |  |  | 17 | NI 9208 | Tri-Loop E BURST
Feedback Signals |  |  | 7 | 4-20 mA | LHS | Water Valve % Closed |  | % | 0 | 100 | 0.004 | 0.02 |  | 7 | 15 | [5 empty] | NI 9208
|  |  |  | K-type TC | Bottom | Vent Temperature [F] |  | F | [12 empty] | NI 9213
|  |  |  | K-type TC | Bottom | Supply Temperature [F] |  | F | [12 empty] | NI 9213
|  |  |  | K-type TC | Bottom | Regulated Temperature [F] |  | F | [12 empty] | NI 9213
|  |  |  | K-type TC | Bottom | Compressor Temperature [F] |  | F | [12 empty] | NI 9213
|  |  |  | K-type TC | Bottom | KAPL Heater Lower RHS [F] |  | F | [12 empty] | NI 9213
|  |  |  | K-type TC | Bottom | KAPL Heater Lower LHS [F] |  | F | [12 empty] | NI 9213
|  |  |  | K-type TC | Bottom | KAPL Heater Upper RHS [F] |  | F | [12 empty] | NI 9213
|  |  |  | K-type TC | Bottom | KAPL Heater Upper LHS [F] |  | F | [12 empty] | NI 9213
|  |  |  | K-type TC | Bottom | Solex HXer? |  | F | [12 empty] | NI 9213
|  |  |  | K-type TC | Bottom | Solex HXer? |  | F | [12 empty] | NI 9213
|  |  |  | K-type TC | Bottom | Solex HXer? |  | F | [12 empty] | NI 9213
|  |  |  | K-type TC | Bottom | Solex HXer? |  | F | [12 empty] | NI 9213
|  |  |  | K-type TC | Bottom | Solex HXer? |  | F | [12 empty] | NI 9213
|  |  |  | K-type TC | Bottom | Solex HXer? |  | F | [12 empty] | NI 9213
|  |  |  | K-type TC | Bottom | Solex HXer? |  | F | [12 empty] | NI 9213
|  |  |  | K-type TC | Bottom | Solex HXer? |  | F | [12 empty] | NI 9213
Hot RTDs | Conax C8334-2, RTD43W8-316SS25-T4-8.00" |  |  | 4-wire RTD | RHS | Mixing Assembly Outlet Temperature |  | F |  |  |  |  | RTD to NI 9219 |  |  |  | 5 | 3 | 4 | 6 | NI 9219 | Change to NI 9216
HXer RTDs | Conax C8334-2, RTD43W8-316SS25-T4-8.00" |  |  | 4-wire RTD | RHS | Solex HXer Bank 4 Outlet Temperature A |  | F |  |  |  |  | RTD to NI 9219 |  |  |  | 5 | 3 | 4 | 6 | NI 9219 | Change to NI 9216
HXer RTDs | Conax C8334-2, RTD43W8-316SS25-T4-8.00" |  |  | 4-wire RTD | RHS | Solex HXer Bank 4 Outlet Temperature B |  | F |  |  |  |  | RTD to NI 9219 |  |  |  | 5 | 3 | 4 | 6 | NI 9219 | Change to NI 9216
HXer RTDs | Conax C8334-2, RTD43W8-316SS25-T4-8.00" |  |  | 4-wire RTD | RHS | Solex HXer Bank 3 Outlet Temperature A |  | F |  |  |  |  | RTD to NI 9219 |  |  |  | 5 | 3 | 4 | 6 | NI 9219 | Change to NI 9216
HXer RTDs | Conax C8334-2, RTD43W8-316SS25-T4-8.00" |  |  | 4-wire RTD | RHS | Solex HXer Bank 3 Outlet Temperature B |  | F |  |  |  |  | RTD to NI 9219 |  |  |  | 5 | 3 | 4 | 6 | NI 9219 | Change to NI 9216
HXer RTDs | Conax C8334-2, RTD43W8-316SS25-T4-8.00" |  |  | 4-wire RTD | RHS | Solex HXer Bank 2 Outlet Temperature A |  | F |  |  |  |  | RTD to NI 9219 |  |  |  | 5 | 3 | 4 | 6 | NI 9219 | Change to NI 9216
HXer RTDs | Conax C8334-2, RTD43W8-316SS25-T4-8.00" |  |  | 4-wire RTD | RHS | Solex HXer Bank 2 Outlet Temperature B |  | F |  |  |  |  | RTD to NI 9219 |  |  |  | 5 | 3 | 4 | 6 | NI 9219 | Change to NI 9216
|  |  |  | 4-wire RTD | RHS |  |  | F |  |  |  |  | RTD to NI 9219 |  |  |  | 5 | 3 | 4 | 6 | NI 9219 | Change to NI 9216

## Signals DEV
Group | Field Component Vender and Part Number | Label (UUID) | Units | Field Signal | Field Wiring Diagram | Cable | Terminal Location | DIN Number | DAQ Signal | Module | DAQ Wiring Diagram
Densities | Emerson CMFS100P999N2BAEKZX_22757 | Pump Outlet Density | kg/m^3 | N/A | Coriolis Flowmeter v2 | 4-pin, 1-keyway, A-coded M12 Patch Cable | LHS | 16 | 4-20 mA Analog Input | NI 9208 | Coriolis to NI 9208 Adapter v1
Densities | Emerson CMFS100P999N2BAEKZX_22757 | Recuperator Cold-side Inlet Density | kg/m^3 | N/A | Coriolis Flowmeter v2 | 4-pin, 1-keyway, A-coded M12 Patch Cable | LHS | 18 | 4-20 mA Analog Input | NI 9208 | Coriolis to NI 9208 Adapter v1
Flows | Emerson CMFS100P999N2BAEKZX_22757 | Pump Mass Flow Rate | kg/s | N/A | Coriolis Flowmeter v2 | 4-pin, 1-keyway, A-coded M12 Patch Cable | LHS | 15 | 4-20 mA Analog Input | NI 9208 | Coriolis to NI 9208 Adapter v1
Flows | Emerson CMFS100P999N2BAEKZX_22757 | Solex HXer Mass Flow Rate | kg/s | N/A | Coriolis Flowmeter v2 | 4-pin, 1-keyway, A-coded M12 Patch Cable | LHS | 17 | 4-20 mA Analog Input | NI 9208 | Coriolis to NI 9208 Adapter v1
Feedback Signals |  | Heater Power Feedback | W | 4-20 mA | KAPL Heater Controller v1 | 4-pin, 1-keyway, A-coded M12 Patch Cable | LHS | 14 | 4-20 mA Analog Input | NI 9208
Pressure Drops | Kobold PAD-HEE5S2NS00 | Pump Inlet Filter Pressure Drop | psid | N/A | Kobold v1 | 4-pin, 1-keyway, A-coded M12 Patch Cable | LHS | 19 | 4-20 mA Analog Input | NI 9208 | Kobold to NI 9208 Adapter v1
Pressure Drops | Kobold PAD-HEE5S2NS00 | Pump Outlet Filter Pressure Drop | psid | N/A | Kobold v1 | 4-pin, 1-keyway, A-coded M12 Patch Cable | LHS | 20 | 4-20 mA Analog Input | NI 9208 | Kobold to NI 9208 Adapter v1
Feedback Signals |  | Pump Frequency Feedback | Hz | 4-20 mA | Magnepump v3 | 4-pin, 1-keyway, A-coded M12 Patch Cable | LHS | 9 | 4-20 mA Analog Input | NI 9208
Pressure Drops | Rosemount 3051SMV3P14A4R3A11A1JL2Q4GMA1039 | Pump Pressure Drop | psid | N/A | N/A | 4-pin, 1-keyway, A-coded M12 Patch Cable | LHS | Tri-Loop A #1 | 4-20 mA Analog Input | NI 9208 | Rosemount Tri-loop to NI 9208 Adapter v1
Pressure Drops | Rosemount | Water Cooler CO2 Pressure Drop | psid | N/A | N/A | 4-pin, 1-keyway, A-coded M12 Patch Cable | LHS | Tri-Loop B #1 | 4-20 mA Analog Input | NI 9208 | Rosemount Tri-loop to NI 9208 Adapter v1
Pressure Drops | Rosemount 3051SMV3P14A4R3A11A1JL2Q4GMA1039 | Recuperator Cold-side Pressure Drop | psid | N/A | N/A | 4-pin, 1-keyway, A-coded M12 Patch Cable | LHS | Tri-Loop C #1 | 4-20 mA Analog Input | NI 9208 | Rosemount Tri-loop to NI 9208 Adapter v1
Pressure Drops | Rosemount 3051SMV3P14A4R3A11A1JL2Q4GMA1039 | Recuperator Hot-side Pressure Drop | psid | N/A | N/A | 4-pin, 1-keyway, A-coded M12 Patch Cable | RHS | Tri-Loop D #1 | 4-20 mA Analog Input | NI 9208 | Rosemount Tri-loop to NI 9208 Adapter v1
Pressure Drops | Rosemount | KAPL Heater Pressure Drop | psid | N/A | N/A | 4-pin, 1-keyway, A-coded M12 Patch Cable | RHS | Tri-Loop E #1 | 4-20 mA Analog Input | NI 9208 | Rosemount Tri-loop to NI 9208 Adapter v1
Pressure Drops | Rosemount 3051SMV3P14A4R3A11A1JL2Q4GMA1039 | Solex HXer Pressure Drop | psid | N/A | N/A | 4-pin, 1-keyway, A-coded M12 Patch Cable | RHS | Tri-Loop F #1 | 4-20 mA Analog Input | NI 9208 | Rosemount Tri-loop to NI 9208 Adapter v1
Pressures | Rosemount 3051SMV3P14A4R3A11A1JL2Q4GMA1039 | Pump Outlet Pressure | psi | N/A | N/A | 4-pin, 1-keyway, A-coded M12 Patch Cable | LHS | Tri-Loop A BURST | 4-20 mA Analog Input | NI 9208 | Rosemount Tri-loop to NI 9208 Adapter v1
Pressures | Rosemount | Water Cooler CO2 Inlet Pressure | psi | N/A | N/A | 4-pin, 1-keyway, A-coded M12 Patch Cable | LHS | Tri-Loop B BURST | 4-20 mA Analog Input | NI 9208 | Rosemount Tri-loop to NI 9208 Adapter v1
Pressures | Rosemount 3051SMV3P14A4R3A11A1JL2Q4GMA1039 | Recuperator Cold-side Inlet Pressure | psi | N/A | N/A | 4-pin, 1-keyway, A-coded M12 Patch Cable | LHS | Tri-Loop C BURST | 4-20 mA Analog Input | NI 9208 | Rosemount Tri-loop to NI 9208 Adapter v1
Pressures | Rosemount 3051SMV3P14A4R3A11A1JL2Q4GMA1039 | Recuperator Hot-side Inlet Pressure | psi | N/A | N/A | 4-pin, 1-keyway, A-coded M12 Patch Cable | RHS | Tri-Loop D BURST | 4-20 mA Analog Input | NI 9208 | Rosemount Tri-loop to NI 9208 Adapter v1
Pressures | Rosemount | KAPL Heater Inlet Pressure | psi | N/A | N/A | 4-pin, 1-keyway, A-coded M12 Patch Cable | RHS | Tri-Loop E BURST | 4-20 mA Analog Input | NI 9208 | Rosemount Tri-loop to NI 9208 Adapter v1
Pressures | Rosemount 3051SMV3P14A4R3A11A1JL2Q4GMA1039 | Solex HXer Inlet Pressure | psi | N/A | N/A | 4-pin, 1-keyway, A-coded M12 Patch Cable | RHS | Tri-Loop F BURST | 4-20 mA Analog Input | NI 9208 | Rosemount Tri-loop to NI 9208 Adapter v1
Water Pressures |  | Water Inlet Pressure | psi | N/A | N/A | 4-pin, 1-keyway, A-coded M12 Patch Cable | LHS | 22 | 4-20 mA Analog Input | NI 9208 | Water Pressure to NI 9208 v1
Water Pressures |  | Water Outlet Pressure | psi | N/A | N/A | 4-pin, 1-keyway, A-coded M12 Patch Cable | LHS | 23 | 4-20 mA Analog Input | NI 9208 | Water Pressure to NI 9208 v1
Cold RTDs | Conax C8334-2, RTD43W8-316SS25-T4-8.00" | Pump Inlet Temperature | F | 4-wire RTD | Rosemount Multivariable Transmitter v3 | 4-pin, 1-keyway, A-coded M12 Patch Cable | LHS | Tri-Loop A #2 | 4-20 mA Analog Input | NI 9208 | Rosemount Tri-loop to NI 9208 Adapter v1
Cold RTDs | Omega PR-22-3-100-A-1/8-0600-M12 | Water Cooler CO2 Outlet Temperature | F | 4-wire RTD | Rosemount Multivariable Transmitter v3 | 4-pin, 1-keyway, A-coded M12 Patch Cable | LHS | Tri-Loop B #2 | 4-20 mA Analog Input | NI 9208 | Rosemount Tri-loop to NI 9208 Adapter v1
Cold RTDs | Omega PR-22-3-100-A-1/8-0600-M12 | Recuperator Cold-side Inlet Temperature | F | 4-wire RTD | Rosemount Multivariable Transmitter v3 | 4-pin, 1-keyway, A-coded M12 Patch Cable | LHS | Tri-Loop C #2 | 4-20 mA Analog Input | NI 9208 | Rosemount Tri-loop to NI 9208 Adapter v1
Hot RTDs | Conax C16791-1, RTD44W4-INC12-T4(HT)-8.00" | Recuperator Cold-side Outlet Temperature | F | 4-wire RTD | Rosemount Multivariable Transmitter v3 | 4-pin, 1-keyway, A-coded M12 Patch Cable | RHS | Tri-Loop D #2 | 4-20 mA Analog Input | NI 9208 | Rosemount Tri-loop to NI 9208 Adapter v1
Hot RTDs | Conax C16791-1, RTD44W4-INC12-T4(HT)-8.00" | KAPL Heater Inlet Temperature | F | 4-wire RTD | Rosemount Multivariable Transmitter v3 | 4-pin, 1-keyway, A-coded M12 Patch Cable | RHS | Tri-Loop E #2 | 4-20 mA Analog Input | NI 9208 | Rosemount Tri-loop to NI 9208 Adapter v1
Hot RTDs | Conax C17049-1, RTD44W4-INC25-T4(HT)-8.00" | Solex HXer Outlet Temperature | F | 4-wire RTD | Rosemount Multivariable Transmitter v3 | 4-pin, 1-keyway, A-coded M12 Patch Cable | RHS | Tri-Loop F #2 | 4-20 mA Analog Input | NI 9208 | Rosemount Tri-loop to NI 9208 Adapter v1
Feedback Signals |  | Recirculation Valve % Closed Feedback | Hz | 4-20 mA | TRIAC Valve Wiring v2 | 4-pin, 1-keyway, A-coded M12 Patch Cable | LHS | 10 | 4-20 mA Analog Input | NI 9208
Feedback Signals |  | Throttle Valve % Closed Feedback | % | 4-20 mA | TRIAC Valve Wiring v2 | 4-pin, 1-keyway, A-coded M12 Patch Cable | LHS | 11 | 4-20 mA Analog Input | NI 9208
Feedback Signals |  | Mixing Coolant Valve % Closed Feedback | % | 4-20 mA | TRIAC Valve Wiring v2 | 4-pin, 1-keyway, A-coded M12 Patch Cable | LHS | 12 | 4-20 mA Analog Input | NI 9208
Water Flow |  | Water Volumetric Flow Rate | gpm | N/A | Water Flowmeter v1 | 4-pin, 1-keyway, A-coded M12 Patch Cable | LHS | 21 | 4-20 mA Analog Input | NI 9208 | Water Flowmeter to NI 9208 v1
Feedback Signals |  | Water Valve % Closed Feedback | % | 4-20 mA | Water Valve v1 | 4-pin, 1-keyway, A-coded M12 Patch Cable | LHS | 13 | 4-20 mA Analog Input | NI 9208
| WIKA | Filling Supply Pressure | psi | 4-20 mA |  | 4-pin, 1-keyway, A-coded M12 Patch Cable | LHS | 24 | 4-20 mA Analog Input | NI 9208
| WIKA | Filling Supply Regulated Pressure | psi | 4-20 mA |  | 4-pin, 1-keyway, A-coded M12 Patch Cable | LHS | 25 | 4-20 mA Analog Input | NI 9208
| WIKA | Compressor Inlet Pressure | psi | 4-20 mA |  | 4-pin, 1-keyway, A-coded M12 Patch Cable | LHS | 26 | 4-20 mA Analog Input | NI 9208
|  | Filling Supply Vent Line Temperature | F | K-type Thermocouple |  | K-type extension wire | Bottom | 1 | K-type Thermocouple | NI 9213
|  | Main Vent Line Temperature | F | K-type Thermocouple |  | K-type extension wire | Bottom | 2 | K-type Thermocouple | NI 9213
|  | Filling Supply Line Temperature | F | K-type Thermocouple |  | K-type extension wire | Bottom | 3 | K-type Thermocouple | NI 9213
|  | Filling Supply Regulated Line Temperature | F | K-type Thermocouple |  | K-type extension wire | Bottom | 4 | K-type Thermocouple | NI 9213
|  | Compressor Inlet Line Temperature | F | K-type Thermocouple |  | K-type extension wire | Bottom | 5 | K-type Thermocouple | NI 9213
|  | Compressor Outlet Line Temperature | F | K-type Thermocouple |  | K-type extension wire | Bottom | 6 | K-type Thermocouple | NI 9213
|  | Rupture Disk Line Temperature | F | K-type Thermocouple |  | K-type extension wire | Bottom | 7 | K-type Thermocouple | NI 9213
|  |  | F | K-type Thermocouple |  | K-type extension wire | Bottom | 8 | K-type Thermocouple | NI 9213
|  |  | F | K-type Thermocouple |  | K-type extension wire | Bottom | 9 | K-type Thermocouple | NI 9213
|  |  | F | K-type Thermocouple |  | K-type extension wire | Bottom | 10 | K-type Thermocouple | NI 9213
|  |  | F | K-type Thermocouple |  | K-type extension wire | Bottom | 11 | K-type Thermocouple | NI 9213
|  |  | F | K-type Thermocouple |  | K-type extension wire | Bottom | 12 | K-type Thermocouple | NI 9213
|  |  | F | K-type Thermocouple |  | K-type extension wire | Bottom | 13 | K-type Thermocouple | NI 9213
|  |  | F | K-type Thermocouple |  | K-type extension wire | Bottom | 14 | K-type Thermocouple | NI 9213
|  |  | F | K-type Thermocouple |  | K-type extension wire | Bottom | 15 | K-type Thermocouple | NI 9213
|  |  | F | K-type Thermocouple |  | K-type extension wire | Bottom | 16 | K-type Thermocouple | NI 9213
Cold RTDs | Conax C8334-2, RTD43W8-316SS25-T4-8.00" | Pump Outlet Temperature | F | 4-wire RTD | N/A | 4-pin, 1-keyway, A-coded M12 Patch Cable | RHS | 5,6 | 4-wire RTD | NI 9216 | 4-wire RTD to Universal 6-pin CB v1
Cold RTDs | Omega PR-22-3-100-A-1/8-0600-M12 | Water Cooler CO2 Inlet Temperature | F | 4-wire RTD | N/A | 4-pin, 1-keyway, A-coded M12 Patch Cable | RHS | 7,8 | 4-wire RTD | NI 9216 | 4-wire RTD to Universal 6-pin CB v1
Cold RTDs | Omega PR-22-3-100-A-1/8-0600-M12 | Recuperator Hot-side Outlet Temperature | F | 4-wire RTD | N/A | 4-pin, 1-keyway, A-coded M12 Patch Cable | RHS | 9,10 | 4-wire RTD | NI 9216 | 4-wire RTD to Universal 6-pin CB v1
Hot RTDs | Conax C16791-1, RTD44W4-INC12-T4(HT)-8.00" | Recuperator Hot-side Inlet Temperature | F | 4-wire RTD | N/A | 4-pin, 1-keyway, A-coded M12 Patch Cable | RHS | 11,12 | 4-wire RTD | NI 9216 | 4-wire RTD to Universal 6-pin CB v1
Hot RTDs | Conax C16791-1, RTD44W4-INC12-T4(HT)-8.00" | KAPL Heater Outlet Temperature | F | 4-wire RTD | N/A | 4-pin, 1-keyway, A-coded M12 Patch Cable | RHS | 13,14 | 4-wire RTD | NI 9216 | 4-wire RTD to Universal 6-pin CB v1
Hot RTDs | Conax C8334-2, RTD43W8-316SS25-T4-8.00" | Solex HXer Inlet Temperature | F | 4-wire RTD | N/A | 4-pin, 1-keyway, A-coded M12 Patch Cable | RHS | 15,16 | 4-wire RTD | NI 9216 | 4-wire RTD to Universal 6-pin CB v1
Hot RTDs | Conax C8334-2, RTD43W8-316SS25-T4-8.00" | Mixing Assembly Outlet Temperature | F | 4-wire RTD | N/A | 4-pin, 1-keyway, A-coded M12 Patch Cable | RHS | 17,18 | 4-wire RTD | NI 9216 | 4-wire RTD to Universal 6-pin CB v1
HXer RTDs | Conax C8334-2, RTD43W8-316SS25-T4-8.00" | Solex HXer Bank 4 Outlet Temperature A | F | 4-wire RTD | N/A | 4-pin, 1-keyway, A-coded M12 Patch Cable | RHS | 19,20 | 4-wire RTD | NI 9216 | 4-wire RTD to Universal 6-pin CB v1
HXer RTDs | Conax C8334-2, RTD43W8-316SS25-T4-8.00" | Solex HXer Bank 4 Outlet Temperature B | F | 4-wire RTD | N/A | 4-pin, 1-keyway, A-coded M12 Patch Cable | RHS | 21,22 | 4-wire RTD | NI 9216 | 4-wire RTD to Universal 6-pin CB v1
HXer RTDs | Conax C8334-2, RTD43W8-316SS25-T4-8.00" | Solex HXer Bank 3 Outlet Temperature A | F | 4-wire RTD | N/A | 4-pin, 1-keyway, A-coded M12 Patch Cable | RHS | 23,24 | 4-wire RTD | NI 9216 | 4-wire RTD to Universal 6-pin CB v1
HXer RTDs | Conax C8334-2, RTD43W8-316SS25-T4-8.00" | Solex HXer Bank 3 Outlet Temperature B | F | 4-wire RTD | N/A | 4-pin, 1-keyway, A-coded M12 Patch Cable | RHS | 25,26 | 4-wire RTD | NI 9216 | 4-wire RTD to Universal 6-pin CB v1
HXer RTDs | Conax C8334-2, RTD43W8-316SS25-T4-8.00" | Solex HXer Bank 2 Outlet Temperature A | F | 4-wire RTD | N/A | 4-pin, 1-keyway, A-coded M12 Patch Cable | RHS | 27,28 | 4-wire RTD | NI 9216 | 4-wire RTD to Universal 6-pin CB v1
HXer RTDs | Conax C8334-2, RTD43W8-316SS25-T4-8.00" | Solex HXer Bank 2 Outlet Temperature B | F | 4-wire RTD | N/A | 4-pin, 1-keyway, A-coded M12 Patch Cable | RHS | 29,30 | 4-wire RTD | NI 9216 | 4-wire RTD to Universal 6-pin CB v1
Water RTDs | Omega PR-22-3-100-A-1/8-0600-M12 | Water Cooler Water Inlet Temperature | F | 4-wire RTD | N/A | 4-pin, 1-keyway, A-coded M12 Patch Cable | RHS | 1,2 | 4-wire RTD | NI 9216 | 4-wire RTD to Universal 6-pin CB v1
Water RTDs | Omega PR-22-3-100-A-1/8-0600-M12 | Water Cooler Water Outlet Temperature | F | 4-wire RTD | N/A | 4-pin, 1-keyway, A-coded M12 Patch Cable | RHS | 3,4 | 4-wire RTD | NI 9216 | 4-wire RTD to Universal 6-pin CB v1
|  |  | F | 4-wire RTD | N/A | 4-pin, 1-keyway, A-coded M12 Patch Cable | RHS | 31,32 | 4-wire RTD | NI 9216 | 4-wire RTD to Universal 6-pin CB v1
mA Outputs |  | Heater Power Command | W | 4-20 mA | KAPL Heater Controller v1 | 4-pin, 1-keyway, A-coded M12 Patch Cable | LHS | 6 | 4-20 mA Analog Output | NI 9266
mA Outputs |  | Pump Frequency Command | Hz | 4-20 mA | Magnepump v3 | 4-pin, 1-keyway, A-coded M12 Patch Cable | LHS | 1 | 4-20 mA Analog Output | NI 9266
mA Outputs |  | Recirculation Valve % Closed Command | Hz | 4-20 mA | TRIAC Valve Wiring v2 | 4-pin, 1-keyway, A-coded M12 Patch Cable | LHS | 2 | 4-20 mA Analog Output | NI 9266
mA Outputs |  | Throttle Valve % Closed Command | % | 4-20 mA | TRIAC Valve Wiring v2 | 4-pin, 1-keyway, A-coded M12 Patch Cable | LHS | 3 | 4-20 mA Analog Output | NI 9266
mA Outputs |  | Mixing Coolant Valve % Closed Command | % | 4-20 mA | TRIAC Valve Wiring v2 | 4-pin, 1-keyway, A-coded M12 Patch Cable | LHS | 4 | 4-20 mA Analog Output | NI 9266
mA Outputs |  | Water Valve % Closed Command | % | 4-20 mA | Water Valve v1 | 4-pin, 1-keyway, A-coded M12 Patch Cable | LHS | 5 | 4-20 mA Analog Output | NI 9266
|  | Drive Air Regulator Command | psi | 4-20 mA |  | 4-pin, 1-keyway, A-coded M12 Patch Cable | LHS | 7 | 4-20 mA Analog Output | NI 9266
|  | Filling Supply Pressure Regulator Command | psi | 4-20 mA |  | 4-pin, 1-keyway, A-coded M12 Patch Cable | LHS | 8 | 4-20 mA Analog Output | NI 9266
Relay Outputs |  | Pump Run Enable Switch | Boolean | 24 VDC | Magnepump v3 | 2-pin, 2-keyway, A-coded M__ Patch Cable | LHS | 27 | Relay | NI 9485 | Run Enable Switch Relay v1
Relay Outputs |  | DAQ Alarm Status | Boolean | 24 VDC | Stack Light v1 | 2-pin, 2-keyway, A-coded M__ Patch Cable | LHS | 28 | Relay | NI 9485 | DAQ Alarm Relay v1
|  | Drive Air Isolation Valve | Boolean | 24 VDC |  | 2-pin, 2-keyway, A-coded M__ Patch Cable | LHS | 29 | Relay | NI 9485
|  | Main Vent Valve Relay | Boolean | 24 VDC |  | 2-pin, 2-keyway, A-coded M__ Patch Cable | LHS | 30 | Relay | NI 9485
|  | Filling Supply Vent Valve Relay | Boolean | 24 VDC |  | 2-pin, 2-keyway, A-coded M__ Patch Cable | LHS | 31 | Relay | NI 9485
|  | Filling Supply Isolation Valve Relay | Boolean | 24 VDC |  | 2-pin, 2-keyway, A-coded M__ Patch Cable | LHS | 32 | Relay | NI 9485

## Uncertainties
Measurement Type | Specification | Uncertainty | LRL | Operating | URL | Units | Vendor | Part Number | Description | Reference
Temperature | RTD44W4-INC12-T4(HT)-8.00" | 0.3 | -250 | 700 | 800 | ⁰C | Conax | C16791-1 | RTD44, Class B Tolerance, 4 Wire Single, Inconel 600 Sheath Material, 0.125" Sheath Diameter, T4 High Temp Encapsulant Terminal Head, and 8" length | Conax RTD Assemblies – Components and Descriptions | 4 (recup ht in, heater in/out, recup lt out)
Temperature | RTD43W8-316SS25-T4-8.00" | 0.3 | -250 | 650 | 600 | ⁰C | Conax | C8334-2 | RTD43, Class B Tolerance, 4 Wire Dual, 316 Stainless Steel Sheath Material, 0.25" Sheath Diameter, T4 Terminal Head, and 8" length | Conax RTD Assemblies – Components and Descriptions | 6 (4x hxer, 2x pump)
Temperature | RTD44W4-INC25-T4(HT)-8.00" | 0.3 | -250 | 550 | 800 | ⁰C | Conax | C17049-1 | RTD44, Class B Tolerance, 4 Wire Single, Inconel 600 Sheath Material, 0.25" Sheath Diameter, T4 High Temp Encapsulant Terminal Head, and 8" length | Conax RTD Assemblies – Components and Descriptions | 4 (2x hxer, hxer oultet, after mixing assembly)
Temperature | PR-22-3-100-A-1/8-0600-M12 | 0.75 | -30 |  | 350 | ⁰C | Omega | PR-22-3-100-A-1/8-0600-M12 | RTD 1/8X6" M12CONN, Class A, 0.00385 TCR | https://www.omega.com/Temperature/pdf/RTDSpecs_Ref.pdf | 7 (2x recup cold side, 4x water cooler, mixing in)
Temperature | 3051SMV3P14A4R3A11A1JL2Q4GMA1039 | .025% of span |  |  |  |  | Rosemount
Temperature | [6 empty] | Rosemount
Pressure | 3051SMV3P14A4R3A11A1JL2Q4GMA1039 | 0.9065 |  |  | 3626 | psi | Rosemount | 6.2484
Pressure |  | 0.9065 |  |  | 3626 | psi | Rosemount | 6.2484
Differential Pressure | PAD-HEE5S2NS00 | 0.02025 | 0.27 |  | 27 | psid | Kobold | PAD-HEE5S2NS00 | Differential Pressure TXMTR 750" WC, 1/4" NPT, Max WP 4499 PSI 4-20MA + HART Cal Span 27 PSI Stainless Steel Same as supplied on SER# 3150233
Differential Pressure | 3051SMV3P14A4R3A11A1JL2Q4GMA1039 | 0.06 |  |  | 150 | psid | Rosemount | 0.41357
Differential Pressure |  | 0.4 |  |  | 1000 | inH2O | Rosemount
Mass Flow | CMFS100P999N2BAEKZX_22757 | 0.0025 |  | 1 |  |  | Emerson
Density | CMFS100P999N2BAEKZX_22757 | 0.5 | 200 | 800 | 5000 | kg/m^3 | Emerson
Volumetric Flow Rate | 1" 5-50 gpm Turbine Flowmeter | 0.15 | 5 | 15 | 50 | gpm | Flomec | GPI/FLOMEC G2-S-10-N-09-GM-A
Pressure |  | 30 | 0 | 1400 | 3000 | psig | WIKA | WIKA A-10